## Dashboard
Coverage Summary
Risk rows extracted | 147
Product requirements | 98
Authoritative PR-to-risk links | 394
Protocol PR coverage rows | 79
Protocol risk references | 344
Copy/reconcile AC rows | 54
Open issues | 107
Navigation
Explorer
Copy to Protocols
Issues
Source Index
Control
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Explorer
Project | RiskScope | RiskSource | RiskLine | RiskSummary | RiskVerification | ProductRequirement | RequirementText | Protocol | OverarchingAcceptanceCriteria | ProtocolAcceptanceCriteria | AcceptanceCriteriaStatus
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-002 | Antibody is not specific and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-023 | The product is used on a non-BOND-III instrument | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-030.1 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-001 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-001 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-001 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-001 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-001 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-001 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-002 | Antibody is not specific and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-028 | The product Interpretation Guide does not explain the scoring system or contain images of expected HER2 expressing breast carcinoma tissue sections | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-039 | The product does not maintain comparable and acceptable performance when scored by different pathologists | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-040 | The product does not maintain comparable and acceptable performance to the same pathologist over time | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-005 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-005 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-009 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-009 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-009 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-011 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-011 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-030.1 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-050 | The product does not maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-053 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-011 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-002 | Antibody is not specific and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-015 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-015 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-017 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-017 | Tissue preparation is not carried out appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-019 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-019 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | Verified against supplied workbook row | PR.10 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim. | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-019.5 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-023 | The product is used on a non-BOND-III instrument | Verified against supplied workbook row | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | DP52-VEP-013 | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | Verified against supplied workbook row | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | DP52-VEP-013 | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-019.5 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | Verified against supplied workbook row | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | DP52-VEP-013 | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | Verified against supplied workbook row | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | DP52-VEP-013 | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-020 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-019 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-020 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-020 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-019 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-020 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-021 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-021 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-021 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-021 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-022.1 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-022.1 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-022.1 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | Verified against supplied workbook row | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | DP52-VEP-013 | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-070 | The product does not have a barcode to allow the BOND System to recognise the identity of the antibody package by means of the on-board scanner | Verified against supplied workbook row | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | DP52-VEP-013 | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-070.5 | The product does not have a barcode to allow the BOND System to recognise the identity of the reagent package by means of the hand held scanner | Verified against supplied workbook row | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | DP52-VEP-013 | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | Verified against supplied workbook row | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-023 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-023 | Scoring of tissue slides results in incorrect result | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | Verified against supplied workbook row | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Verified against supplied workbook row | PR.03 | The product, when used as part of the HER2 BMA CDx assay system, shall be capable of confirming correct execution of the assay when used in accordance with recommended assay parameters. |  | Deviations in the assay reagents, instrument and protocol that alter the HER2 expression status in patient samples should result in a failure of the CLC to meet the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Verified against supplied workbook row | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | DP61-VEP-002 | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Verified against supplied workbook row | PR.07 | The Product labelling must describe the expected staining patterns and characteristics of each cell cores. | DP61-VEP-002 | HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to confirm they describes the expected staining patterns and characteristics of each cell line core. Sample Size: TBD | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-039 | The product does not maintain comparable and acceptable performance when scored by different pathologists | Verified against supplied workbook row | PR.23 | The Product, when used within the HER2 BMA CDx Assay System, must generate staining that can be interpreted in the same way by different pathologists (inter-reader precision). |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-040 | The product does not maintain comparable and acceptable performance to the same pathologist over time | Verified against supplied workbook row | PR.24 | The Product, when used within the HER2 BMA CDx Assay System, must generate staining that can be interpreted in the same way by the same pathologist over time (intra-reader precision). |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Verified against supplied workbook row | PR.25 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of HER2 BMA Antibody CDx on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. This acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%.Point estimate of ≥90% when the lower bound of the 95% Confident Interval is ≥85%. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-030.1 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.26 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | Verified against supplied workbook row | PR.27 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different BOND-III systems. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | Verified against supplied workbook row | PR.28 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different SSAs on a BOND-III Instrument run batch. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC
DP52 | No linked risk |  |  |  | No risk citation in DP52-PR-001 | PR.29 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining within BOND Runs (intra-run repeatability). | DP52-VEP-007 | For each of the conditions listed below, a minimum of 90% of tissue cases must demonstrate Acceptable Staining and Consistent Staining across the same SSA on a BOND-III Instrument. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | Verified against supplied workbook row | PR.30 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | DP52-VEP-007 | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | Verified against supplied workbook row | PR.31 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining in different laboratories (inter-site reproducibility). |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | No linked risk |  |  |  | No risk citation in DP52-PR-001 | PR.32 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (Dako HercepTest ™ ) for HER2 overexpression. |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | No linked risk |  |  |  | No risk citation in DP52-PR-001 | PR.33 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (VENTANA® PATHWAY® anti-HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody) for HER2-low and HER2-ultralow. |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | No linked risk |  |  |  | No risk citation in DP52-PR-001 | PR.34 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (VENTANA® PATHWAY® anti-HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody) for HER2 overexpression. |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | Verified against supplied workbook row | PR.35 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | DP52/DP61-VEP-001 (Environmental) | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96 °C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-024.5 | CLC slide is not evaluated as intended | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-088 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | Verified against supplied workbook row | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. |  | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating temperature range of 18-26°C (64-79°F on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-032.5 | Shipping packaging does not protects the product during transport | Verified against supplied workbook row | PR.49 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | Assessment results in 100% samples demonstrating that packaging of products contained within the different shipping boxes being tested (ZSV003, ZCB033, ZCB292, ZCB172, ZCB035, and ZCB088) is undamaged and serviceable via visual inspection when the shipping box is: • dropped from a minimum height of 0.8m 2 , onto a hard surface, on: o the face of the packaging with the largest surface area o one corner of the shipping package pointing downwards • Boxes are stacked following EU regulations along with other box types and shipments (max height of 150cm). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-089 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | Verified against supplied workbook row | PR.37 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III humidity specifications. |  | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating humidity (noncondensing) 30 to 80% RH range on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | Verified against supplied workbook row | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | DP52-VEP-003 | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-088 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | Verified against supplied workbook row | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | DP52-VEP-003 | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-089 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | Verified against supplied workbook row | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | DP52-VEP-003 | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-056 | The product is not able to run on the BOND-III System at the same time as other antibodies and probes | Verified against supplied workbook row | PR.39 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | DP52-VEP-013 | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance.The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. Stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-057 | The product is not able to run on the BOND-III System at the same time as other antibodies and probes | Verified against supplied workbook row | PR.39 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | DP52-VEP-013 | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance.The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. Stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | COPY/RECONCILE OVERARCHING AC
DP52 | No linked risk |  |  |  | No risk citation in DP52-PR-001 | PR.40 | The Product can be run within the same BOND Run Batch as other BOND Detection systems and staining protocol combinations. | DP52-VEP-013 | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-033.5 | Outer packaging does not protect the product during opening | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-033.5 | Outer packaging does not protect the product during opening | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-072 | The product labels do not have font and symbols that is legible to laboratory users | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-073 | The product labels do not have font, symbols and barcode that is legible to laboratory users | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-074 | The product labels do not have font and symbols that is legible to laboratory users | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-075 | The product labels do not have font and symbols that is legible to laboratory users | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-039 | Reagent spills during opening of the reagent container | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-039 | Reagent spills during opening of the reagent container | Verified against supplied workbook row | PR.47 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging remains serviceable following 46 thermal cycles designed to mimic on BOND and storage conditions (i.e. 7hr at maximum expected temperature of 34ºC, storage at 2-8ºC). AND One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-042 | The product (antibody) is used beyond the expiry date | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-078 | The product container label text and graphics are not human readable after the product is exposed to product, BOND or detection kit reagents | Verified against supplied workbook row | PR.44 | The symbols and text included on The Product labelling must be human legible. | DP52-VEP-003 | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-081 | The product packaging labels are not legible after exposure to normal storage and handling conditions | Verified against supplied workbook row | PR.44 | The symbols and text included on The Product labelling must be human legible. | DP52-VEP-003 | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | Verified against supplied workbook row | PR.45 | The Product when expired, will not allow a staining run to commence. | DP52-VEP-013 | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument.Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-042 | The product (antibody) is used beyond the expiry date | Verified against supplied workbook row | PR.45 | The Product when expired, will not allow a staining run to commence. | DP52-VEP-013 | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument.Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-068 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-042 | The product (antibody) is used beyond the expiry date | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-043 | The product (CLC) is used beyond the expiry date | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-043 | The product (CLC) is used beyond the expiry date | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life when stored in the immediate packaging (minimum of six months shelf life with a minimum two months safety margin). | DP61-VEP-005 | Following a minimum of 14 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. The Product must also retains its ability to correctly detect procedural failures. Sample Size: TBD | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.Following a minimum of 18 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The Product must also retain its ability to correctly detect procedural failures. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | Verified against supplied workbook row | PR.47 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging remains serviceable following 46 thermal cycles designed to mimic on BOND and storage conditions (i.e. 7hr at maximum expected temperature of 34ºC, storage at 2-8ºC). AND One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-085 | The product does not withstand and/or leaks, after transportation to the customer | Verified against supplied workbook row | PR.47 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging remains serviceable following 46 thermal cycles designed to mimic on BOND and storage conditions (i.e. 7hr at maximum expected temperature of 34ºC, storage at 2-8ºC). AND One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-085 | The product does not withstand and/or leaks, after transportation to the customer | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-090 | Outer packaging (pouch) does not remain intact during transit to customer | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-091 | Outer packaging (pouch) does not remain intact during transit to customer | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-046 | User does not dispose of reagents appropriately | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-092 | Shipping packaging does not remain intact during transit to customer | Verified against supplied workbook row | PR.49 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | Assessment results in 100% samples demonstrating that packaging of products contained within the different shipping boxes being tested (ZSV003, ZCB033, ZCB292, ZCB172, ZCB035, and ZCB088) is undamaged and serviceable via visual inspection when the shipping box is: • dropped from a minimum height of 0.8m 2 , onto a hard surface, on: o the face of the packaging with the largest surface area o one corner of the shipping package pointing downwards • Boxes are stacked following EU regulations along with other box types and shipments (max height of 150cm). |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-046 | User does not dispose of reagents appropriately | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-080 | The product container label text and graphics are not BOND readable after the product is exposed to product, BOND or detection kit reagents | Verified against supplied workbook row | PR.50 | The barcodes included on The Product reagent container label must remain fully readable by the BOND-III fully automated system following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-078 | The product container label text and graphics are not human readable after the product is exposed to product, BOND or detection kit reagents | Verified against supplied workbook row | PR.51 | The barcodes, symbols and text included on The Product reagent container label must remain human legible following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-081 | The product packaging labels are not legible after exposure to normal storage and handling conditions | Verified against supplied workbook row | PR.51 | The barcodes, symbols and text included on The Product reagent container label must remain human legible following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-082 | The product packaging labels are not legible after exposure to normal storage and handling conditions | Verified against supplied workbook row | PR.51 | The barcodes, symbols and text included on The Product reagent container label must remain human legible following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | No linked risk |  |  |  | No risk citation in DP52-PR-001 | PR.52 | The Product shall include compliance with UDI specifications. | DP52-VEP-003 | The labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | The final labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.53 | The Product labelling must be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | DP52-VEP-003 | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.54 | The Product labelling must be compliant with the relevant requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | DP52-VEP-003 | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.55 | The Product labelling must be compliant with the requirements of a Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | DP52-VEP-003 | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.56 | The Product labelling must be compliant with the relevant UK general IVD per the UK MDR 2002 labelling regulations. | DP52-VEP-003 | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.57 | The Product must only be compatible with BOND-III fully automated systems operating BOND Software that is compliant with regional regulatory requirements. |  | The Product can only be run on BOND software that is compliant with regional regulatory requirements. AND The Product Specifications should require the reagent container barcode to meet the DHFR 2119 BOND Barcode Specification. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | No linked risk |  |  |  | No risk citation in DP52-PR-001 | PR.58 | The Product, when used within the HER2 BMA CDx Assay System, is expected to show greater agreement with reflex testing results, helping enable a more definitive determination of HER2 status based on the initial IHC test | DP52.DP61-VAP-002 | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-030.1 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-050 | The product does not maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-053 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-046 | User does not dispose of reagents appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-048 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-048 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-048 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-049 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-049 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-049 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-050 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-050 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-028 | The product Interpretation Guide does not explain the scoring system or contain images of expected HER2 expressing breast carcinoma tissue sections | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-039 | The product does not maintain comparable and acceptable performance when scored by different pathologists | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-040 | The product does not maintain comparable and acceptable performance to the same pathologist over time | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-050 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-052 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-052 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-052 | The BOND run is not set up correctly | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-053 | The user coverslips using an non-appropriate mountant | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-053 | The user coverslips using an non-appropriate mountant | Verified against supplied workbook row | PR.64 | The Product must be compatible with reagent‑grade alcohol, xylene, and xylene or toluene‑based mounting media. | DP52-VEP-008 | All tissue cases must maintain acceptable staining when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA CDx Assay System, all slides must maintain acceptable staining when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA CDx Assay System. | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | COPY/RECONCILE OVERARCHING AC
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-053 | The user coverslips using an non-appropriate mountant | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-019 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-023 | The product is used on a non-BOND-III instrument | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-028 | The product Interpretation Guide does not explain the scoring system or contain images of expected HER2 expressing breast carcinoma tissue sections | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-030.1 |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-050 | The product does not maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-053 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-068 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-087 | The product is not compatible with clearing procedures using reagent grade alcohol, xylene, and a xylene/toluene-based mounting media | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-088 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-089 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-055 | The product is used for staining on aged FFPE tissue sections | Verified against supplied workbook row | PR.17 | The Product will have a minimum cut section stability of 30 days. |  | 100% of all HER2 test slides stored and aged at 2-8°C stained with The Product must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated cut section stability claim. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-055 | The product is used for staining on aged FFPE tissue sections | Verified against supplied workbook row | PR.18 | The Product labelling will detail known limitations of formalin fixed, paraffin-embedded (FFPE) cut section stability. | DP52-VEP-003 | Product Instructions for Use (IFU) document specifies any known limitations for cut tissue sections during the stability period as determined by verification testing. | Product Instructions For Use (IFU) document specifies the recommended storage conditions for cut tissue sections during the stability period (temperature, humidity, protection from light). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-055 | The product is used for staining on aged FFPE tissue sections | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-058 | The BOND-III instrument is not maintained as per IFU and BOND User Manual | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-058 | The BOND-III instrument is not maintained as per IFU and BOND User Manual | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-058 | The BOND-III instrument is not maintained as per IFU and BOND User Manual | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Verified against supplied workbook row | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-068 | The product is not used for automated staining on the BOND-III staining platform | Verified against supplied workbook row | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | DP52-VEP-013 | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-068 | The product is not used for automated staining on the BOND-III staining platform | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-068 | The product is not used for automated staining on the BOND-III staining platform | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-070 | The product is ran on the BOND-III using a non recommended BOND software version | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-070 | The product is ran on the BOND-III using a non recommended BOND software version | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-072 | Assay is not run alongside CLCs | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-072 | Assay is not run alongside CLCs | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-072 | Assay is not run alongside CLCs | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-075 | Product is not run alongside recommended tissue controls | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-075 | Product is not run alongside recommended tissue controls | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-075 | Product is not run alongside recommended tissue controls | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | Verified against supplied workbook row | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | Verified against supplied workbook row | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. |  | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating temperature range of 18-26°C (64-79°F on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-068 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-081 | The product packaging labels are not legible after exposure to normal storage and handling conditions | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-086 | Slides are not dehydrated and cleared as per standard IHC techniques | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | DP52-specific design risk | LBN-dRMW-016 | dRMW-risk-087 | The product is not compatible with clearing procedures using reagent grade alcohol, xylene, and a xylene/toluene-based mounting media | Verified against supplied workbook row | PR.64 | The Product must be compatible with reagent‑grade alcohol, xylene, and xylene or toluene‑based mounting media. | DP52-VEP-008 | All tissue cases must maintain acceptable staining when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA CDx Assay System, all slides must maintain acceptable staining when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA CDx Assay System. | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | COPY/RECONCILE OVERARCHING AC
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-086 | Slides are not dehydrated and cleared as per standard IHC techniques | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-007 | The product does not contain three cell line cores which demonstrate a staining range (high expression, low expression and negative) | Verified against supplied workbook row | PR.02 | The Product must contain cell cores representative of different HER2 expression levels (null HER2 expression, low HER2 expression and high HER2 expression). |  | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-009 | The product is not compatible with components of the HER2 BMA Assay System | Verified against supplied workbook row | PR.02 | The Product must contain cell cores representative of different HER2 expression levels (null HER2 expression, low HER2 expression and high HER2 expression). |  | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-010 | The Cell Line Control (CLC) does not function as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | Verified against supplied workbook row | PR.02 | The Product must contain cell cores representative of different HER2 expression levels (null HER2 expression, low HER2 expression and high HER2 expression). |  | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-010 | The Cell Line Control (CLC) does not function as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | Verified against supplied workbook row | PR.03 | The product, when used as part of the HER2 BMA CDx assay system, shall be capable of confirming correct execution of the assay when used in accordance with recommended assay parameters. |  | Deviations in the assay reagents, instrument and protocol that alter the HER2 expression status in patient samples should result in a failure of the CLC to meet the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | Verified against supplied workbook row | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-009 | The product is not compatible with components of the HER2 BMA Assay System | Verified against supplied workbook row | PR.04 | The Product must be compatible for use with all components of the HER2 BMA CDx Assay System. | DP61-VEP-006 | The Product, when stained using all components of the HER2 BMA Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained using all components of the HER2 BMA CDx Assay system, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using all components of the HER2 BMA CDx Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-013 | The product IFU and interpretation guide does not explain the scoring system or contain images of expected stained cell lines | Verified against supplied workbook row | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | DP61-VEP-002 | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-040 | The product is not able to be evaluated consistently by different pathologists | Verified against supplied workbook row | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | DP61-VEP-002 | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-013 | The product IFU and interpretation guide does not explain the scoring system or contain images of expected stained cell lines | Verified against supplied workbook row | PR.07 | The Product labelling must describe the expected staining patterns and characteristics of each cell cores. | DP61-VEP-002 | HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to confirm they describes the expected staining patterns and characteristics of each cell line core. Sample Size: TBD | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-018 | The product does not maintain acceptable staining performance after transportation to the customer | Verified against supplied workbook row | PR.08 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | DP61-VEP-005 | The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | 100% of stained Product must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide after being subject to transport simulation.The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-001 | The product is not ready to use and requires processing prior to use | Verified against supplied workbook row | PR.09 | The Product must generate Acceptable Staining without having to be subjected to any immediate preparation (e.g., slide baking or incubation) after being removed from the recommended storage conditions (2-8Â°C). | DP61-VEP-006 | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-004 | The product cannot be used directly from recommended storage | Verified against supplied workbook row | PR.09 | The Product must generate Acceptable Staining without having to be subjected to any immediate preparation (e.g., slide baking or incubation) after being removed from the recommended storage conditions (2-8Â°C). | DP61-VEP-006 | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP61 | No linked risk |  |  |  | No risk citation in DP61-PR-001 | PR.10 | The Product must generate Acceptable Staining when used in conjunction with the recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols. | DP61-VEP-006 | The Product, when stained using the HER2 BMA Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-017 | Each cell line does not maintain its HER2 expression level over the shelf life when stored at recommended temperature, and does not have a minimum shelf life of 6 months for initial release | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life when stored in the immediate packaging (minimum of six months shelf life with a minimum two months safety margin). | DP61-VEP-005 | Following a minimum of 14 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. The Product must also retains its ability to correctly detect procedural failures. Sample Size: TBD | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.Following a minimum of 18 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The Product must also retain its ability to correctly detect procedural failures. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-039 | The product is not reproducible and does not have consistent appearance/ morphology between lots | Verified against supplied workbook row | PR.12 | The Intermediate Product must be stable for a minimum of six months when stored at room temperature. | DP61-VEP-005 | The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide after being stored at room temperature for at least six months. Sample Size: TBD | Following a minimum of six months storage at room temperature, 100% of The Intermediate Product (individual agarose cell line cores of BT474-M3, RT4 and PANC-1) continues to produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1 must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide after being stored at room temperature for at least six months. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-021 | The product does not maintain HER2 expression levels when different lots of the CLC are used | Verified against supplied workbook row | PR.13 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | DP61-VEP-003 | Three lots of the Product, manufactured from separate multiblocks will be stained using a single lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by one trained pathologist, with the lot of the Product as the only variable. Staining results must align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must demonstrate Acceptable Staining and Consistent Staining across six different lots. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.For inter-lot testing, the lower bound of the two-sided 95% confidence interval for the overall percent agreement should be at least 85%. Testing will be performed on 6 lots. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-039 | The product is not reproducible and does not have consistent appearance/ morphology between lots | Verified against supplied workbook row | PR.13 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | DP61-VEP-003 | Three lots of the Product, manufactured from separate multiblocks will be stained using a single lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by one trained pathologist, with the lot of the Product as the only variable. Staining results must align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must demonstrate Acceptable Staining and Consistent Staining across six different lots. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.For inter-lot testing, the lower bound of the two-sided 95% confidence interval for the overall percent agreement should be at least 85%. Testing will be performed on 6 lots. | COPY/RECONCILE OVERARCHING AC
DP61 | No linked risk |  |  |  | No risk citation in DP61-PR-001 | PR.14 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining throughout the same lots of The Product (intra-lot repeatability). | DP61-VEP-003 | Sections taken from the beginning, middle, and end of a single Cell Line Control multiblock will be stained using one lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. A trained pathologist will evaluate the slides, with the section position within the multiblock as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs For intra-lot testing, taken from within a single lot must generate Acceptable Staining and Consistent Staining. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate the overall percent agreement should be at least 85%. Only sections within a single lot will be tested. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-022 | The product does not maintain HER2 expression levels when stained on multiple BOND-III instruments over multiple days | Verified against supplied workbook row | PR.15 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained on five non-consecutive days (spanning a minimum period of 20 days). A single trained pathologist will score the slides, with the staining day being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained over 5 non-consecutive days (spanning a minimum period of 20 days), CLCs must generate Acceptable Staining and Consistent Staining. For inter-day testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate for the overall percent agreement should should be at least 85%. Testing will occur over 5 non-consecutive days spanning a minimum period of 20 days. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-022 | The product does not maintain HER2 expression levels when stained on multiple BOND-III instruments over multiple days | Verified against supplied workbook row | PR.16 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained using three different BOND-III fully automated systems. A single trained pathologist will score the slides, with the BOND-III instrument being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across three different BOND-IIIs. For inter-instrument testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed across three BOND-IIIs. | COPY/RECONCILE OVERARCHING AC
DP61 | No linked risk |  |  |  | No risk citation in DP61-PR-001 | PR.17 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained on different BOND Runs/Slide Staining Assemblies (SSAs) within the same BOND Run Batch. A single trained pathologist will score the slides, with the BOND Run/SSA as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across different SSAs on a BOND-III Instrument run batch and within BOND Runs (SSAs). For inter-run testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Iner-run data will be collected across BOND run batches. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-021.2 | The product does not maintain HER2 expression levels when different lots of the detection kit DS9800 are used | Verified against supplied workbook row | PR.18 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Polymer Refine Detection (DS9800) system. The slides will be scored by a single trained pathologist, with the lot of the DS9800 as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system For inter-detectiontested samples lot repeatability testing, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Three lots will be used during testing. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-021.1 | The product does not maintain HER2 expression levels when different lots of the antibody are used | Verified against supplied workbook row | PR.19 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx are varied (inter-antibody lot repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by a single trained pathologist, with the lot of the antibody as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. For inter-antibody lot repeatability testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed using 3 lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-047 | The product does not maintain HER2 expression levels across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | Verified against supplied workbook row | PR.20 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. |  | Cell Line Control slides from BOND runs performed at environmental temperatures between 18-26°C must stain as expected - Core A (BT474-M3) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-048 | The product does not maintain HER2 expression levels across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | Verified against supplied workbook row | PR.21 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III humidity specifications. |  | Cell Line Control slides from BOND runs performed at environmental humidities between 30-80% Relative Humidity must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-006 | The FFPE cell cores are not presented in every section in the appropriate orientation | Verified against supplied workbook row | PR.22 | The placement of the cell line cores within The Product must be consistently orientated relative to each other. | DP61-VEP-006 | HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA Assay Interpretation Guide. Sample Size: TBD | A HER2 scoring SME/qualified pathologist will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product IFU and HER2 BMA CDx Assay Interpretation Guide. 100% of assessed CLC slides must demonstrate correct core orientation. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-028 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | Verified against supplied workbook row | PR.23 | The Product must have labelling that provides all required hazard and safety information and associated symbols. | DP61-VEP-002 | Users will review the Product slide mailer label, hazard notification slip and IFU and complete a checklist to confirm that the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents) is present. Sample Size: TBD | Users will review the Product slide mailer label, IFU and product specifications and complete a checklist to confirm that the required safety information is present. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-028 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | Verified against supplied workbook row | PR.24 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP61-VEP-002 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. Sample Size: TBD | Laboratory users will confirm that the hazards detailed in the IFU and product specifications can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP). | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-032 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-032.5 | The outer and immediate packaging does not remain intact during shipping | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-033 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-033.5 | The outer and immediate packaging does not remain intact during shipping | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-032 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-032.5 | The outer and immediate packaging does not remain intact during shipping | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-033 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-033.5 | The outer and immediate packaging does not remain intact during shipping | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-032.5 | The outer and immediate packaging does not remain intact during shipping | Verified against supplied workbook row | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | DP61-VEP-002 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-050 | The product container (slide mailer) does not remain intact during shipment to customer | Verified against supplied workbook row | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | DP61-VEP-002 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-051 | The product container (slide mailer) is not safe to handle | Verified against supplied workbook row | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | DP61-VEP-002 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-031 | The product and immediate container labels containing product, lot number and expiry date are not legible | Verified against supplied workbook row | PR.28 | The on-slide printed text must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-041 | The product labels do not have font and symbols that are human legible | Verified against supplied workbook row | PR.28 | The on-slide printed text must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-043 | The product labels do not have font and symbols that is legible to laboratory users | Verified against supplied workbook row | PR.29 | The Immediate Packaging label must be capable of withstanding shipping and storage within the recommended temperature range and normal use without becoming detached. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remain fully adhered and undamaged. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remained fully adhered and undamaged. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-044 | The product container labels do not remain adhered during the entire life of the product | Verified against supplied workbook row | PR.29 | The Immediate Packaging label must be capable of withstanding shipping and storage within the recommended temperature range and normal use without becoming detached. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remain fully adhered and undamaged. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remained fully adhered and undamaged. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-041 | The product labels do not have font and symbols that are human legible | Verified against supplied workbook row | PR.30 | The symbols and text included on The Product labelling must be human legible. | DP61-VEP-002 | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC
DP61 | No linked risk |  |  |  | No risk citation in DP61-PR-001 | PR.31 | The Product shall include compliance with UDI specifications. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.32 | The Product labelling must be created to be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.33 | The Product labelling must be created to be compliant with the relevant requirements from IVDR EU/2017/746. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.34 | The Product labelling must be created to be compliant with the relevant Canadian IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | Verified against supplied workbook row | PR.35 | The Product labelling must be created to comply with the relevant UK general IVD as per the UK MDR 2002 labelling regulations. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-049 | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND III user manual, 0-1600m | Verified against supplied workbook row | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | DP52/DP61-VEP-001 (Environmental) | Cell Line Control slides from BOND runs performed at altitudes up to and including 1600 meters above sea level must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBC | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-004 | The product cannot be used directly from recommended storage | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-006 | The FFPE cell cores are not presented in every section in the appropriate orientation | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-007 | The product does not contain three cell line cores which demonstrate a staining range (high expression, low expression and negative) | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-010 | The Cell Line Control (CLC) does not function as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-011 | The Cell Line Control (CLC) functions as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-013 | The product IFU and interpretation guide does not explain the scoring system or contain images of expected stained cell lines | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-028 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-033 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-033.5 | The outer and immediate packaging does not remain intact during shipping | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-049 | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND III user manual, 0-1600m | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-049.01 | The product does not maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-049.1 | The product does not maintain HER2 expression levels and morphology when used above 1600m | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-050 | The product container (slide mailer) does not remain intact during shipment to customer | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-052 | The product is not ran with the recommended HIER protocol | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-092 | The product is not ready to use and requires mixing or shaking before use | Verified against supplied workbook row | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | DP52-VEP-013 | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-092 | The product is not ready to use and requires mixing or shaking before use | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-093 | Use of the delayed start function for extended periods of time when running the HER2 BMA CDx Assay | Verified against supplied workbook row | PR.18 | The Product labelling will detail known limitations of formalin fixed, paraffin-embedded (FFPE) cut section stability. | DP52-VEP-003 | Product Instructions for Use (IFU) document specifies any known limitations for cut tissue sections during the stability period as determined by verification testing. | Product Instructions For Use (IFU) document specifies the recommended storage conditions for cut tissue sections during the stability period (temperature, humidity, protection from light). | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-093 | Use of the delayed start function for extended periods of time when running the HER2 BMA CDx Assay | Verified against supplied workbook row | PR.41 | The Product must generate Acceptable Staining when there is a delay of up to seven days between placing The Product on the BOND-III and commencement of staining. | DP52-VEP-008 | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-093 | Use of the delayed start function for extended periods of time when running the HER2 BMA CDx Assay | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-096 | The BOND run completes successfully and the user does not unload the SSA in a timely manner | Verified against supplied workbook row | PR.42 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining when there is a delay of up to 24 hours between completion of staining and removal. | DP52-VEP-008 | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-096 | The BOND run completes successfully and the user does not unload the SSA in a timely manner | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-096 | The BOND run completes successfully and the user does not unload the SSA in a timely manner | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Verified against supplied workbook row | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Verified against supplied workbook row | PR.35 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | DP52/DP61-VEP-001 (Environmental) | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96 °C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC
DP52 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING
DP61 | Shared DP52/DP61 end-user risk | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-011 | The Cell Line Control (CLC) functions as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | Verified against supplied workbook row | PR.38 | The Product must be compatible with reagent grade alcohol, xylene and toluene based mounting media. |  | All Cell Line Controls must maintain acceptable staining and cell morphology when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA Assay System. All Cell Line Controls must maintain acceptable staining and cell morphology when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA Assay System. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-049.01 | The product does not maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude | Verified exact ID exists | N/A | The product shall maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude | DP52/DP61-VEP-001 (Environmental) | N/A - risk-characterisation-only protocol row | 100% of sections stained at 1600m using the alternative HIER protocols (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins) should demonstrate no evidence of altitude staining artifact impacting slide interpretation or altered HER2 status as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | N/A - no PR reference in protocol
DP61 | DP61-specific design risk | LBN-dRMW-017 | dRMW-risk-049.1 | The product does not maintain HER2 expression levels and morphology when used above 1600m | Verified exact ID exists | N/A | The product shall maintain HER2 expression levels and morphology when used above 1600m | DP52/DP61-VEP-001 (Environmental) | N/A - risk-characterisation-only protocol row | 100% of sections stained at 1800m should demonstrate no evidence of altitude staining artifact impacting slide interpretation or altered HER2 status as defined in Section 5.3.1. The HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. If 100% of sections do not stain satisfactorily, then either the positive/negative tissue controls or CLCs must fail. | N/A - no PR reference in protocol

## Control
HER2 Risk / Requirement / Protocol Traceability Matrix
Document status | DRAFT - requires qualified QA review before controlled use
Document ID | [NEEDS INPUT]
Revision | [NEEDS INPUT]
Generated | 2026-09-22
Purpose | Visualize DP52 and DP61 as separate projects, each with its own design risks, while preserving their shared links to LBN-euRMW-020 end-user risks.
Source reconciliation | DP61 risk-to-PR links now use signed DP61-PR-001 Rev 05. Any PRs present in the overarching plan but absent from the signed PR are listed in Issues.
Use | Start on Dashboard. Filter by Project and Risk Scope first, then narrow by risk line, PR, protocol, or acceptance-criteria status. Shared euRMW lines may correctly appear under both projects.
Verification basis | Risk IDs were matched to exact supplied workbook rows. Source identifier variants and missing citations are listed in Issues.

## Traceability (hidden)
ProductFamily | RiskSourceDocument | RiskID | RiskHazard | RiskHazardousSituation | RiskWorkbookRow | RiskVerification | ProductRequirementID | ProductRequirementText | ProtocolID | OverarchingAcceptanceCriteria | ProtocolAcceptanceCriteria | AcceptanceCriteriaCopyStatus | LinkBasis
DP52 | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | 8 | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-002 | Antibody is not specific and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | 9 | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | 10 | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | 11 | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 13 | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.01 | The Product must be specific for the HER2 protein. | DP52-VEP-011 | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-023 | The product is used on a non-BOND-III instrument | The product is loaded onto a non- BOND-III instrument and accepted | 20 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | The product is loaded onto a BOND III instrument and is accepted and ran on a non-validated version of BOND software | 21 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | The product stains inappropriately using current BOND-III bulk reagents and ancillaries | 22 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-030.1 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | The product stains sub-optimally using the recommended staining protocol | 26 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-015 | Tissue preparation is not carried out appropriately | User mounts tissue section outside of the recommended slide staining area, resulting in sub-optimal staining | 16 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-017 | Tissue preparation is not carried out appropriately | User mounts tissue section onto non-recommended glass slides resulting in sub-optimal staining | 17 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-048 | The BOND run is not set up correctly | Covertiles misplaced on the slide leading to sub-optimal staining | 35 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-049 | The BOND run is not set up correctly | Covertiles are not placed on top of the slides leading to sub-optimal staining | 36 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-050 | The BOND run is not set up correctly | Unclean covertiles are used for staining leading to background staining and/or sub-optimal staining | 37 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-052 | The BOND run is not set up correctly | Insufficient BOND ancillary reagents loaded onto BOND leading to BOND run error or delay | 38 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-058 | The BOND-III instrument is not maintained as per IFU and BOND User Manual | The BOND-III instrument is not cleaned and mould/contaminants in tubing lead to sub-optimal staining. | 41 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Slide is set up as 100 µL dispense and the user attempts to use it on a BOND-III instrument | 42 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Slides stained using alternative detection kit to the recommended DS9800 leading to sub-optimal staining | 43 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Negative reagent controls are not included in run and sub-optimal staining cannot be detected | 49 | Verified against supplied workbook row | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | DP52-VEP-013 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | 8 | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-002 | Antibody is not specific and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | 9 | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | 10 | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | 11 | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 13 | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | DP52-VEP-005 | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 13 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-028 | The product Interpretation Guide does not explain the scoring system or contain images of expected HER2 expressing breast carcinoma tissue sections | The Interpretation Guide is inadequate to facilitate HER2 scoring | 23 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-039 | The product does not maintain comparable and acceptable performance when scored by different pathologists | Staining of the tissue is interpretated differently by different pathologists | 27 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-040 | The product does not maintain comparable and acceptable performance to the same pathologist over time | Staining of the tissue is interpretated differently by the same pathologist on a different day leading to inconsistent diagnosis | 28 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | User prepares sections from non-breast carcinoma tissue resulting in sub-optimal staining | 15 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-019 | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2 overexpression/HER2-low/HER2-ultralow case as HER2-null | 18 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-019.5 | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2 overexpression as a HER2-low or HER2-ultralow | 19 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-020 | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2-null as HER2-low/HER2-ultralow/HER2 overexpression (including scoring artefactual/non-specific staining) | 20 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-021 | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain not visible leading to error in scoring | 21 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-022.1 | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain too strong leading to error in scoring | 22 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-023 | Scoring of tissue slides results in incorrect result | Pathologist scores using an alternative scoring method resulting in incorrect result | 23 | Verified against supplied workbook row | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | DP52-VEP-003 | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | 10 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Insufficient volume dispensed on slide leading to poor staining quality | 17 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | The product stains inappropriately using current BOND-III bulk reagents and ancillaries | 22 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Tissue sections are stained sub-optimally when different lots of antibody are used | 24 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-030.1 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | The product results in sub-optimal staining when used on different BOND III instruments | 29 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | The product is loaded onto the BOND and results in sub-optimal staining between different SSAs within the same BOND-III instrument | 30 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | The product is loaded onto the BOND-III and results in sub-optimal staining when used on a different day | 31 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-050 | The product does not maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | Staining becomes sub-optimal over the life of the reagent due to dilution of reagents | 32 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-053 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Staining becomes sub-optimal over the life of the reagent due to evaporation | 33 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 36 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | The product does not have the minimum stability | 37 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 38 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 39 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | The product requires mixing or shaking to produce optimal staining | 40 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | The product is loaded onto the BOND III instrument for 16 & 23 runs and is not stable for all cycles of use | 42 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | User bakes the CLC resulting in sub-optimal staining | 14 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Incorrect interpretation of the cell line controls may result in a run being unnecessarily voided | 24 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-072 | Assay is not run alongside CLCs | CLCs are not included in run and sub-optimal staining cannot be detected CLCs are not included in run and certain changes in procedural staining performance may not be identified | 46 | Verified against supplied workbook row | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | DP52-VEP-008 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | 8 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-002 | Antibody is not specific and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | 9 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-003 | Antibody is not sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | 10 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | 11 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 13 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-001 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | User embeds tissue specimens with a non-standard paraffin/embedding reagent. The tissue sections are stained sub-optimally | 8 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-005 | Tissue preparation is not carried out appropriately | User prepares tissue sections outside that of the recommended thickness. The tissue sections are stained sub-optimally | 9 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-009 | Tissue preparation is not carried out appropriately | User bakes tissue sections at different temperatures to that recommended resulting in sub-optimal staining | 10 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-011 | Tissue preparation is not carried out appropriately | User bakes tissue sections at different time-frame to that recommended and results in sub-optimal staining | 11 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | User prepares sections from non-breast carcinoma tissue resulting in sub-optimal staining | 15 | Verified against supplied workbook row | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-010 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | 8 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | 11 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-001 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | User embeds tissue specimens with a non-standard paraffin/embedding reagent. The tissue sections are stained sub-optimally | 8 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-009 | Tissue preparation is not carried out appropriately | User bakes tissue sections at different temperatures to that recommended resulting in sub-optimal staining | 10 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-011 | Tissue preparation is not carried out appropriately | User bakes tissue sections at different time-frame to that recommended and results in sub-optimal staining | 11 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | User prepares sections from non-breast carcinoma tissue resulting in sub-optimal staining | 15 | Verified against supplied workbook row | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | DP52-VEP-003 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | 8 | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | 11 | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-001 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 38 | Verified against supplied workbook row | PR.10 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim. | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | The product results in sub-optimal staining when used on different BOND III instruments | 29 | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | The product is loaded onto the BOND and results in sub-optimal staining between different SSAs within the same BOND-III instrument | 30 | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | The product is loaded onto the BOND-III and results in sub-optimal staining when used on a different day | 31 | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | The product is loaded onto the BOND III instrument for 16 & 23 runs and is not stable for all cycles of use | 42 | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | User leaves product on the BOND-III for extended periods of time between use, leading to antibody degradation and evaporation | 53 | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | DP52-VEP-012 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-023 | The product is used on a non-BOND-III instrument | The product is loaded onto a non- BOND-III instrument and accepted | 20 | Verified against supplied workbook row | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | DP52-VEP-013 | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | The product is loaded onto a BOND III instrument and is accepted and ran on a non-validated version of BOND software | 21 | Verified against supplied workbook row | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | DP52-VEP-013 | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-068 | The product is not used for automated staining on the BOND-III staining platform | The product is used on an alternative automated staining platform to the BOND-III leading to sub-optimal staining | 44 | Verified against supplied workbook row | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | DP52-VEP-013 | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 39 | Verified against supplied workbook row | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | DP52-VEP-013 | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | The product requires mixing or shaking to produce optimal staining | 40 | Verified against supplied workbook row | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | DP52-VEP-013 | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-092 | The product is not ready to use and requires mixing or shaking before use | User shakes the product prior to use introducing bubbles, resulting in sub-optimal staining | 54 | Verified against supplied workbook row | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | DP52-VEP-013 | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-019 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained, and the lab does not have additional stock of antibody units | 15 | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-020 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained, and the lab does not have additional stock of antibody units | 16 | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Insufficient volume dispensed on slide leading to poor staining quality | 17 | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Slide is set up as 100 µL dispense and the user attempts to use it on a BOND-III instrument | 42 | Verified against supplied workbook row | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-019 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained, and the lab does not have additional stock of antibody units | 15 | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-020 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained, and the lab does not have additional stock of antibody units | 16 | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Insufficient volume dispensed on slide leading to poor staining quality | 17 | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-021 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained | 18 | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Slide is set up as 100 µL dispense and the user attempts to use it on a BOND-III instrument | 42 | Verified against supplied workbook row | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 36 | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | The product does not have the minimum stability | 37 | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-042 | The product (antibody) is used beyond the expiry date | The product is used beyond the expiry date leading to suboptimal staining | 30 | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-043 | The product (CLC) is used beyond the expiry date | The user uses the product (CLC) beyond the expiry date and suboptimal staining leads to a incorrect void run | 31 | Verified against supplied workbook row | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | DP52-VEP-012 | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-055 | The product is used for staining on aged FFPE tissue sections | FFPE tissue sections used beyond antigen stability leading to sub-optimal staining | 40 | Verified against supplied workbook row | PR.17 | The Product will have a minimum cut section stability of 30 days. |  | 100% of all HER2 test slides stored and aged at 2-8°C stained with The Product must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated cut section stability claim. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-055 | The product is used for staining on aged FFPE tissue sections | FFPE tissue sections used beyond antigen stability leading to sub-optimal staining | 40 | Verified against supplied workbook row | PR.18 | The Product labelling will detail known limitations of formalin fixed, paraffin-embedded (FFPE) cut section stability. | DP52-VEP-003 | Product Instructions for Use (IFU) document specifies any known limitations for cut tissue sections during the stability period as determined by verification testing. | Product Instructions For Use (IFU) document specifies the recommended storage conditions for cut tissue sections during the stability period (temperature, humidity, protection from light). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-093 | Use of the delayed start function for extended periods of time when running the HER2 BMA CDx Assay | User uses BOND-III delayed start function which results in sub-optimal staining | 55 | Verified against supplied workbook row | PR.18 | The Product labelling will detail known limitations of formalin fixed, paraffin-embedded (FFPE) cut section stability. | DP52-VEP-003 | Product Instructions for Use (IFU) document specifies any known limitations for cut tissue sections during the stability period as determined by verification testing. | Product Instructions For Use (IFU) document specifies the recommended storage conditions for cut tissue sections during the stability period (temperature, humidity, protection from light). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | The product is loaded onto a BOND III instrument and is accepted and ran on a non-validated version of BOND software | 21 | Verified against supplied workbook row | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | DP52-VEP-013 | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-070 | The product does not have a barcode to allow the BOND System to recognise the identity of the antibody package by means of the on-board scanner | The product is loaded onto the BOND-III Instrument and the antibody cannot be recognised by the on-board scanner and cannot automatically be run | 43 | Verified against supplied workbook row | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | DP52-VEP-013 | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-070.5 | The product does not have a barcode to allow the BOND System to recognise the identity of the reagent package by means of the hand held scanner | The product cannot be registered into the BOND inventory by scanning the side barcode | 44 | Verified against supplied workbook row | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | DP52-VEP-013 | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | The product stains sub-optimally using the recommended staining protocol | 26 | Verified against supplied workbook row | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Slides stained using alternative detection kit to the recommended DS9800 leading to sub-optimal staining | 43 | Verified against supplied workbook row | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | User selects a non-recommended staining protocol leading to sub-optimal staining | 48 | Verified against supplied workbook row | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not provide acceptable staining across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps | 57 | Verified against supplied workbook row | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-001 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 52 | Verified against supplied workbook row | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Staining in tissue samples differs to the staining that would have resulted if ran at altitudes below 1600m Epitope retrieval failure leads to sub optimal staining or patchy staining | 58 | Verified against supplied workbook row | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | DP52-VEP-013 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | The product stains sub-optimally using the recommended staining protocol | 26 | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | User selects a non-recommended staining protocol leading to sub-optimal staining | 48 | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 52 | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 57 | Verified against supplied workbook row | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | DP52-VEP-003 | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-039 | The product does not maintain comparable and acceptable performance when scored by different pathologists | Staining of the tissue is interpretated differently by different pathologists | 27 | Verified against supplied workbook row | PR.23 | The Product, when used within the HER2 BMA CDx Assay System, must generate staining that can be interpreted in the same way by different pathologists (inter-reader precision). |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-040 | The product does not maintain comparable and acceptable performance to the same pathologist over time | Staining of the tissue is interpretated differently by the same pathologist on a different day leading to inconsistent diagnosis | 28 | Verified against supplied workbook row | PR.24 | The Product, when used within the HER2 BMA CDx Assay System, must generate staining that can be interpreted in the same way by the same pathologist over time (intra-reader precision). |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Tissue sections are stained sub-optimally when different lots of antibody are used | 24 | Verified against supplied workbook row | PR.25 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of HER2 BMA Antibody CDx on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. This acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%.Point estimate of ≥90% when the lower bound of the 95% Confident Interval is ≥85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-030.1 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.26 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | The product results in sub-optimal staining when used on different BOND III instruments | 29 | Verified against supplied workbook row | PR.27 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different BOND-III systems. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | The product is loaded onto the BOND and results in sub-optimal staining between different SSAs within the same BOND-III instrument | 30 | Verified against supplied workbook row | PR.28 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | DP52-VEP-007 | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different SSAs on a BOND-III Instrument run batch. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | [5 empty] | No risk citation in DP52-PR-001 | PR.29 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining within BOND Runs (intra-run repeatability). | DP52-VEP-007 | For each of the conditions listed below, a minimum of 90% of tissue cases must demonstrate Acceptable Staining and Consistent Staining across the same SSA on a BOND-III Instrument. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | The product is loaded onto the BOND-III and results in sub-optimal staining when used on a different day | 31 | Verified against supplied workbook row | PR.30 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | DP52-VEP-007 | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | The product results in sub-optimal staining when used on different BOND III instruments | 29 | Verified against supplied workbook row | PR.31 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining in different laboratories (inter-site reproducibility). |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | [5 empty] | No risk citation in DP52-PR-001 | PR.32 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (Dako HercepTest ™ ) for HER2 overexpression. |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | [5 empty] | No risk citation in DP52-PR-001 | PR.33 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (VENTANA® PATHWAY® anti-HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody) for HER2-low and HER2-ultralow. |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | [5 empty] | No risk citation in DP52-PR-001 | PR.34 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (VENTANA® PATHWAY® anti-HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody) for HER2 overexpression. |  | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not provide acceptable staining across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps | 57 | Verified against supplied workbook row | PR.35 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | DP52/DP61-VEP-001 (Environmental) | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96 °C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Staining in tissue samples differs to the staining that would have resulted if ran at altitudes below 1600m Epitope retrieval failure leads to sub optimal staining or patchy staining | 58 | Verified against supplied workbook row | PR.35 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | DP52/DP61-VEP-001 (Environmental) | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96 °C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-088 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual | 59 | Verified against supplied workbook row | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. |  | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating temperature range of 18-26°C (64-79°F on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | User does not store product at recommended temperature range (2-8°C). Following storage the product is used on the BOND-III and results in sub-optimal staining | 50 | Verified against supplied workbook row | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. |  | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating temperature range of 18-26°C (64-79°F on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-089 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidities defined in the BOND user manual | 60 | Verified against supplied workbook row | PR.37 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III humidity specifications. |  | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating humidity (noncondensing) 30 to 80% RH range on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not provide acceptable staining across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps | 57 | Verified against supplied workbook row | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | DP52-VEP-003 | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-088 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual | 59 | Verified against supplied workbook row | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | DP52-VEP-003 | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-089 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidities defined in the BOND user manual | 60 | Verified against supplied workbook row | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | DP52-VEP-003 | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-056 | The product is not able to run on the BOND-III System at the same time as other antibodies and probes | The product is loaded onto the BOND-III instrument and is unable to schedule testing alongside other antibodies and probes. Laboratory has to wait until run is complete before they can run the HER2 BMA CDx Assay | 34 | Verified against supplied workbook row | PR.39 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | DP52-VEP-013 | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance.The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. Stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-057 | The product is not able to run on the BOND-III System at the same time as other antibodies and probes | The product is loaded onto the BOND-III instrument and it completes testing with BOND errors alongside other antibodies and probes | 35 | Verified against supplied workbook row | PR.39 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | DP52-VEP-013 | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance.The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. Stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | [5 empty] | No risk citation in DP52-PR-001 | PR.40 | The Product can be run within the same BOND Run Batch as other BOND Detection systems and staining protocol combinations. | DP52-VEP-013 | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-093 | Use of the delayed start function for extended periods of time when running the HER2 BMA CDx Assay | User uses BOND-III delayed start function which results in sub-optimal staining | 55 | Verified against supplied workbook row | PR.41 | The Product must generate Acceptable Staining when there is a delay of up to seven days between placing The Product on the BOND-III and commencement of staining. | DP52-VEP-008 | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-096 | The BOND run completes successfully and the user does not unload the SSA in a timely manner | The user does not unload the SSA, the slides is continually hydrated and staining becomes sub-optimal over time | 56 | Verified against supplied workbook row | PR.42 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining when there is a delay of up to 24 hours between completion of staining and removal. | DP52-VEP-008 | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 45 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-072 | The product labels do not have font and symbols that is legible to laboratory users | The product is opened and the user cannot read the labelling font or the symbols on the container labels so cannot identify the product | 46 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-073 | The product labels do not have font, symbols and barcode that is legible to laboratory users | The product is opened and the user cannot read the labelling font, symbols or barcode on the container labels so cannot identify the product | 47 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-074 | The product labels do not have font and symbols that is legible to laboratory users | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | 48 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-075 | The product labels do not have font and symbols that is legible to laboratory users | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | 49 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | The user comes into contact with the reagent which contains 0.035% 2-methylisothiazol-3(2H)-one | 32 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-046 | User does not dispose of reagents appropriately | The user does not dispose of the product appropriately. The component parts (0.035% 2-methylisothiazol-3(2H)-one) are dangerous for the environment | 33 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Glass slides not properly disposed and may cause injury | 34 | Verified against supplied workbook row | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | DP52-VEP-003 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-078 | The product container label text and graphics are not human readable after the product is exposed to product, BOND or detection kit reagents | The user cannot read the labelling font or the symbols on the container labels after exposure to product reagents or BOND reagents, rendering the user unable to identify relevant safety labelling | 51 | Verified against supplied workbook row | PR.44 | The symbols and text included on The Product labelling must be human legible. | DP52-VEP-003 | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-081 | The product packaging labels are not legible after exposure to normal storage and handling conditions | After normal storage and handling the labels are not legible by the user, rendering the user unable to identify the reagent/relevant safety labelling | 53 | Verified against supplied workbook row | PR.44 | The symbols and text included on The Product labelling must be human legible. | DP52-VEP-003 | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 36 | Verified against supplied workbook row | PR.45 | The Product when expired, will not allow a staining run to commence. | DP52-VEP-013 | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument.Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-042 | The product (antibody) is used beyond the expiry date | The product is used beyond the expiry date leading to suboptimal staining | 30 | Verified against supplied workbook row | PR.45 | The Product when expired, will not allow a staining run to commence. | DP52-VEP-013 | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument.Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-068 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by good laboratory practices | 41 | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-033.5 | Outer packaging does not protect the product during opening | Reagent container leaks exposing user to hazardous reagent | 27 | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-039 | Reagent spills during opening of the reagent container | User opens reagent container and is exposed to hazardous reagent | 29 | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | The user comes into contact with the reagent which contains 0.035% 2-methylisothiazol-3(2H)-one | 32 | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-046 | User does not dispose of reagents appropriately | The user does not dispose of the product appropriately. The component parts (0.035% 2-methylisothiazol-3(2H)-one) are dangerous for the environment | 33 | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Glass slides not properly disposed and may cause injury | 34 | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | User leaves product on the BOND-III for extended periods of time between use, leading to antibody degradation and evaporation | 53 | Verified against supplied workbook row | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP52-VEP-003 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | The product leaks after transport, exposing the user to the antibody | 55 | Verified against supplied workbook row | PR.47 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging remains serviceable following 46 thermal cycles designed to mimic on BOND and storage conditions (i.e. 7hr at maximum expected temperature of 34ºC, storage at 2-8ºC). AND One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-085 | The product does not withstand and/or leaks, after transportation to the customer | The product reagents leak after transport, reducing the working volume of the product | 56 | Verified against supplied workbook row | PR.47 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging remains serviceable following 46 thermal cycles designed to mimic on BOND and storage conditions (i.e. 7hr at maximum expected temperature of 34ºC, storage at 2-8ºC). AND One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-039 | Reagent spills during opening of the reagent container | User opens reagent container and is exposed to hazardous reagent | 29 | Verified against supplied workbook row | PR.47 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging remains serviceable following 46 thermal cycles designed to mimic on BOND and storage conditions (i.e. 7hr at maximum expected temperature of 34ºC, storage at 2-8ºC). AND One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | The product leaks after transport, exposing the user to the antibody | 55 | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-085 | The product does not withstand and/or leaks, after transportation to the customer | The product reagents leak after transport, reducing the working volume of the product | 56 | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-090 | Outer packaging (pouch) does not remain intact during transit to customer | Hazard notification slip may fall out and user won't be aware of hazards and know how to find IFU | 61 | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-091 | Outer packaging (pouch) does not remain intact during transit to customer | Hazard notification slip and products may be exposed to environmental condition affecting legibility of text and labels | 62 | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-033.5 | Outer packaging does not protect the product during opening | Reagent container leaks exposing user to hazardous reagent | 27 | Verified against supplied workbook row | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-092 | Shipping packaging does not remain intact during transit to customer | Product containers may fall out and not be provided to user | 63 | Verified against supplied workbook row | PR.49 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | Assessment results in 100% samples demonstrating that packaging of products contained within the different shipping boxes being tested (ZSV003, ZCB033, ZCB292, ZCB172, ZCB035, and ZCB088) is undamaged and serviceable via visual inspection when the shipping box is: • dropped from a minimum height of 0.8m 2 , onto a hard surface, on: o the face of the packaging with the largest surface area o one corner of the shipping package pointing downwards • Boxes are stacked following EU regulations along with other box types and shipments (max height of 150cm). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-032.5 | Shipping packaging does not protects the product during transport | Reagent container leaks during transport, exposing user to hazardous reagent | 26 | Verified against supplied workbook row | PR.49 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. |  | Assessment results in 100% samples demonstrating that packaging of products contained within the different shipping boxes being tested (ZSV003, ZCB033, ZCB292, ZCB172, ZCB035, and ZCB088) is undamaged and serviceable via visual inspection when the shipping box is: • dropped from a minimum height of 0.8m 2 , onto a hard surface, on: o the face of the packaging with the largest surface area o one corner of the shipping package pointing downwards • Boxes are stacked following EU regulations along with other box types and shipments (max height of 150cm). |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-080 | The product container label text and graphics are not BOND readable after the product is exposed to product, BOND or detection kit reagents | The BOND III instrument is unable to read the barcodes of the products and therefore the BOND run cannot schedule | 52 | Verified against supplied workbook row | PR.50 | The barcodes included on The Product reagent container label must remain fully readable by the BOND-III fully automated system following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-078 | The product container label text and graphics are not human readable after the product is exposed to product, BOND or detection kit reagents | The user cannot read the labelling font or the symbols on the container labels after exposure to product reagents or BOND reagents, rendering the user unable to identify relevant safety labelling | 51 | Verified against supplied workbook row | PR.51 | The barcodes, symbols and text included on The Product reagent container label must remain human legible following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-081 | The product packaging labels are not legible after exposure to normal storage and handling conditions | After normal storage and handling the labels are not legible by the user, rendering the user unable to identify the reagent/relevant safety labelling | 53 | Verified against supplied workbook row | PR.51 | The barcodes, symbols and text included on The Product reagent container label must remain human legible following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-082 | The product packaging labels are not legible after exposure to normal storage and handling conditions | After normal storage and handling the labels are not legible by the BOND system | 54 | Verified against supplied workbook row | PR.51 | The barcodes, symbols and text included on The Product reagent container label must remain human legible following routine shipping and storage. |  | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | [5 empty] | No risk citation in DP52-PR-001 | PR.52 | The Product shall include compliance with UDI specifications. | DP52-VEP-003 | The labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | The final labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 45 | Verified against supplied workbook row | PR.53 | The Product labelling must be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | DP52-VEP-003 | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 45 | Verified against supplied workbook row | PR.54 | The Product labelling must be compliant with the relevant requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | DP52-VEP-003 | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 45 | Verified against supplied workbook row | PR.55 | The Product labelling must be compliant with the requirements of a Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | DP52-VEP-003 | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 45 | Verified against supplied workbook row | PR.56 | The Product labelling must be compliant with the relevant UK general IVD per the UK MDR 2002 labelling regulations. | DP52-VEP-003 | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 45 | Verified against supplied workbook row | PR.57 | The Product must only be compatible with BOND-III fully automated systems operating BOND Software that is compliant with regional regulatory requirements. |  | The Product can only be run on BOND software that is compliant with regional regulatory requirements. AND The Product Specifications should require the reagent container barcode to meet the DHFR 2119 BOND Barcode Specification. |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | [5 empty] | No risk citation in DP52-PR-001 | PR.58 | The Product, when used within the HER2 BMA CDx Assay System, is expected to show greater agreement with reflex testing results, helping enable a more definitive determination of HER2 status based on the initial IHC test | DP52.DP61-VAP-002 | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | 8 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | 11 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Insufficient volume dispensed on slide leading to poor staining quality | 17 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | The product stains inappropriately using current BOND-III bulk reagents and ancillaries | 22 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Tissue sections are stained sub-optimally when different lots of antibody are used | 24 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-030.1 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | The product results in sub-optimal staining when used on different BOND III instruments | 29 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | The product is loaded onto the BOND and results in sub-optimal staining between different SSAs within the same BOND-III instrument | 30 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | The product is loaded onto the BOND-III and results in sub-optimal staining when used on a different day | 31 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-050 | The product does not maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | Staining becomes sub-optimal over the life of the reagent due to dilution of reagents | 32 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-053 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Staining becomes sub-optimal over the life of the reagent due to evaporation | 33 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 36 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | The product does not have the minimum stability | 37 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 38 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 39 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | The product requires mixing or shaking to produce optimal staining | 40 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | The product is loaded onto the BOND III instrument for 16 & 23 runs and is not stable for all cycles of use | 42 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-001 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | User embeds tissue specimens with a non-standard paraffin/embedding reagent. The tissue sections are stained sub-optimally | 8 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | User bakes the CLC resulting in sub-optimal staining | 14 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-021 | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain not visible leading to error in scoring | 21 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-022.1 | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain too strong leading to error in scoring | 22 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-042 | The product (antibody) is used beyond the expiry date | The product is used beyond the expiry date leading to suboptimal staining | 30 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-075 | Product is not run alongside recommended tissue controls | Tissue controls are not included in run and sub-optimal staining cannot be detected | 47 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Negative reagent controls are not included in run and sub-optimal staining cannot be detected | 49 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | User does not store product at recommended temperature range (2-8°C). Following storage the product is used on the BOND-III and results in sub-optimal staining | 50 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 52 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | User leaves product on the BOND-III for extended periods of time between use, leading to antibody degradation and evaporation | 53 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 57 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Staining in tissue samples differs to the staining that would have resulted if ran at altitudes below 1600m Epitope retrieval failure leads to sub optimal staining or patchy staining | 58 | Verified against supplied workbook row | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | DP52-VEP-013 | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-011 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 13 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-028 | The product Interpretation Guide does not explain the scoring system or contain images of expected HER2 expressing breast carcinoma tissue sections | The Interpretation Guide is inadequate to facilitate HER2 scoring | 23 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-039 | The product does not maintain comparable and acceptable performance when scored by different pathologists | Staining of the tissue is interpretated differently by different pathologists | 27 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-040 | The product does not maintain comparable and acceptable performance to the same pathologist over time | Staining of the tissue is interpretated differently by the same pathologist on a different day leading to inconsistent diagnosis | 28 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-019 | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2 overexpression/HER2-low/HER2-ultralow case as HER2-null | 18 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-019.5 | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2 overexpression as a HER2-low or HER2-ultralow | 19 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-020 | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2-null as HER2-low/HER2-ultralow/HER2 overexpression (including scoring artefactual/non-specific staining) | 20 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-021 | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain not visible leading to error in scoring | 21 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-022.1 | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain too strong leading to error in scoring | 22 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-023 | Scoring of tissue slides results in incorrect result | Pathologist scores using an alternative scoring method resulting in incorrect result | 23 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Incorrect interpretation of the cell line controls may result in a run being unnecessarily voided | 24 | Verified against supplied workbook row | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | DP52-VEP-003 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-001 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | 8 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-007 | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | 11 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-010 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | 12 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-016 | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | 14 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-019 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained, and the lab does not have additional stock of antibody units | 15 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-020.5 | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Insufficient volume dispensed on slide leading to poor staining quality | 17 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-023 | The product is used on a non-BOND-III instrument | The product is loaded onto a non- BOND-III instrument and accepted | 20 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-024 | The product is used on a BOND software version software version 8.x or greater) | The product is loaded onto a BOND III instrument and is accepted and ran on a non-validated version of BOND software | 21 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-025 | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | The product stains inappropriately using current BOND-III bulk reagents and ancillaries | 22 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-028 | The product Interpretation Guide does not explain the scoring system or contain images of expected HER2 expressing breast carcinoma tissue sections | The Interpretation Guide is inadequate to facilitate HER2 scoring | 23 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-029 | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Tissue sections are stained sub-optimally when different lots of antibody are used | 24 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-030.1 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-034 | The product is not compatible with the recommended BOND staining protocols | The product stains sub-optimally using the recommended staining protocol | 26 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-042 | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | The product results in sub-optimal staining when used on different BOND III instruments | 29 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-043 | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | The product is loaded onto the BOND and results in sub-optimal staining between different SSAs within the same BOND-III instrument | 30 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-046 | The product does not maintain comparable and acceptable performance when used across different days | The product is loaded onto the BOND-III and results in sub-optimal staining when used on a different day | 31 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-050 | The product does not maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | Staining becomes sub-optimal over the life of the reagent due to dilution of reagents | 32 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-053 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Staining becomes sub-optimal over the life of the reagent due to evaporation | 33 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-058 | The product is used after it is no longer stable | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 36 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-059 | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | The product does not have the minimum stability | 37 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-060 | The product does not maintain acceptable staining performance after undergoing transportation to the customer | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 38 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-062 | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | 39 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-066 | The product is not ready to use and requires mixing or shaking before use | The product requires mixing or shaking to produce optimal staining | 40 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-068 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by good laboratory practices | 41 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-069 | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | The product is loaded onto the BOND III instrument for 16 & 23 runs and is not stable for all cycles of use | 42 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | The product leaks after transport, exposing the user to the antibody | 55 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not provide acceptable staining across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps | 57 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-087 | The product is not compatible with clearing procedures using reagent grade alcohol, xylene, and a xylene/toluene-based mounting media | The product is stained on BOND but is not compatible with laboratory clearing and mounting procedures which results in difficulty in interpretation | 58 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-088 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual | 59 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-089 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidities defined in the BOND user manual | 60 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-001 |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-003 | Tissue preparation is not carried out appropriately | User embeds tissue specimens with a non-standard paraffin/embedding reagent. The tissue sections are stained sub-optimally | 8 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-005 | Tissue preparation is not carried out appropriately | User prepares tissue sections outside that of the recommended thickness. The tissue sections are stained sub-optimally | 9 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-009 | Tissue preparation is not carried out appropriately | User bakes tissue sections at different temperatures to that recommended resulting in sub-optimal staining | 10 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-011 | Tissue preparation is not carried out appropriately | User bakes tissue sections at different time-frame to that recommended and results in sub-optimal staining | 11 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | User bakes the CLC resulting in sub-optimal staining | 14 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-013 | Tissue preparation is not carried out appropriately | User prepares sections from non-breast carcinoma tissue resulting in sub-optimal staining | 15 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-015 | Tissue preparation is not carried out appropriately | User mounts tissue section outside of the recommended slide staining area, resulting in sub-optimal staining | 16 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-017 | Tissue preparation is not carried out appropriately | User mounts tissue section onto non-recommended glass slides resulting in sub-optimal staining | 17 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Incorrect interpretation of the cell line controls may result in a run being unnecessarily voided | 24 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | The user comes into contact with the reagent which contains 0.035% 2-methylisothiazol-3(2H)-one | 32 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-046 | User does not dispose of reagents appropriately | The user does not dispose of the product appropriately. The component parts (0.035% 2-methylisothiazol-3(2H)-one) are dangerous for the environment | 33 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Glass slides not properly disposed and may cause injury | 34 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-048 | The BOND run is not set up correctly | Covertiles misplaced on the slide leading to sub-optimal staining | 35 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-049 | The BOND run is not set up correctly | Covertiles are not placed on top of the slides leading to sub-optimal staining | 36 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-050 | The BOND run is not set up correctly | Unclean covertiles are used for staining leading to background staining and/or sub-optimal staining | 37 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-052 | The BOND run is not set up correctly | Insufficient BOND ancillary reagents loaded onto BOND leading to BOND run error or delay | 38 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-053 | The user coverslips using an non-appropriate mountant | User coverslips using non-recommended mountant, leading to sub-optimal visualization of staining | 39 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-055 | The product is used for staining on aged FFPE tissue sections | FFPE tissue sections used beyond antigen stability leading to sub-optimal staining | 40 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-058 | The BOND-III instrument is not maintained as per IFU and BOND User Manual | The BOND-III instrument is not cleaned and mould/contaminants in tubing lead to sub-optimal staining. | 41 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-064 | A dispense volume of 150 µL is not used during application of the product to tissue sections | Slide is set up as 100 µL dispense and the user attempts to use it on a BOND-III instrument | 42 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Slides stained using alternative detection kit to the recommended DS9800 leading to sub-optimal staining | 43 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-068 | The product is not used for automated staining on the BOND-III staining platform | The product is used on an alternative automated staining platform to the BOND-III leading to sub-optimal staining | 44 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-070 | The product is ran on the BOND-III using a non recommended BOND software version | The BOND-III is on a non-recommended software version leading to sub-optimal staining | 45 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-072 | Assay is not run alongside CLCs | CLCs are not included in run and sub-optimal staining cannot be detected CLCs are not included in run and certain changes in procedural staining performance may not be identified | 46 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-075 | Product is not run alongside recommended tissue controls | Tissue controls are not included in run and sub-optimal staining cannot be detected | 47 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | User selects a non-recommended staining protocol leading to sub-optimal staining | 48 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Negative reagent controls are not included in run and sub-optimal staining cannot be detected | 49 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | User does not store product at recommended temperature range (2-8°C). Following storage the product is used on the BOND-III and results in sub-optimal staining | 50 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-086 | Slides are not dehydrated and cleared as per standard IHC techniques | User does not appropriately dehydrate and clear the sections resulting in bubbles/debris on slides that interfere with slide interpretation | 51 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 52 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | User leaves product on the BOND-III for extended periods of time between use, leading to antibody degradation and evaporation | 53 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-092 | The product is not ready to use and requires mixing or shaking before use | User shakes the product prior to use introducing bubbles, resulting in sub-optimal staining | 54 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-093 | Use of the delayed start function for extended periods of time when running the HER2 BMA CDx Assay | User uses BOND-III delayed start function which results in sub-optimal staining | 55 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-096 | The BOND run completes successfully and the user does not unload the SSA in a timely manner | The user does not unload the SSA, the slides is continually hydrated and staining becomes sub-optimal over time | 56 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 57 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Staining in tissue samples differs to the staining that would have resulted if ran at altitudes below 1600m Epitope retrieval failure leads to sub optimal staining or patchy staining | 58 | Verified against supplied workbook row | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP52-VEP-003 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-068 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by good laboratory practices | 41 | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-071 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 45 | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-081 | The product packaging labels are not legible after exposure to normal storage and handling conditions | After normal storage and handling the labels are not legible by the user, rendering the user unable to identify the reagent/relevant safety labelling | 53 | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-084 | The product does not withstand and/or leaks, after transportation to the customer | The product leaks after transport, exposing the user to the antibody | 55 | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-044 | Incorrect /no PPE is worn when using the product | The user comes into contact with the reagent which contains 0.035% 2-methylisothiazol-3(2H)-one | 32 | Verified against supplied workbook row | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | DP52-VEP-003 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-dRMW-016 | dRMW-risk-087 | The product is not compatible with clearing procedures using reagent grade alcohol, xylene, and a xylene/toluene-based mounting media | The product is stained on BOND but is not compatible with laboratory clearing and mounting procedures which results in difficulty in interpretation | 58 | Verified against supplied workbook row | PR.64 | The Product must be compatible with reagent‑grade alcohol, xylene, and xylene or toluene‑based mounting media. | DP52-VEP-008 | All tissue cases must maintain acceptable staining when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA CDx Assay System, all slides must maintain acceptable staining when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA CDx Assay System. | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP52 | LBN-euRMW-020 | euRMW-risk-053 | The user coverslips using an non-appropriate mountant | User coverslips using non-recommended mountant, leading to sub-optimal visualization of staining | 39 | Verified against supplied workbook row | PR.64 | The Product must be compatible with reagent‑grade alcohol, xylene, and xylene or toluene‑based mounting media. | DP52-VEP-008 | All tissue cases must maintain acceptable staining when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA CDx Assay System, all slides must maintain acceptable staining when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA CDx Assay System. | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP52-PR-001 Rev 06; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-007 | The product does not contain three cell line cores which demonstrate a staining range (high expression, low expression and negative) | Product does not contain cores including a range of HER2 expression so is unable to act as a procedural control to detect a failed run | 11 | Verified against supplied workbook row | PR.02 | The Product must contain cell cores representative of different HER2 expression levels (null HER2 expression, low HER2 expression and high HER2 expression). |  | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-009 | The product is not compatible with components of the HER2 BMA Assay System | Product is not compatible with HER2 BMA Assay System and does not stain correctly resulting in a false void run | 12 | Verified against supplied workbook row | PR.02 | The Product must contain cell cores representative of different HER2 expression levels (null HER2 expression, low HER2 expression and high HER2 expression). |  | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-010 | The Cell Line Control (CLC) does not function as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The functionality demonstrated by the CLC is matrix‑dependent and may differ from that observed in tissue controls | 13 | Verified against supplied workbook row | PR.02 | The Product must contain cell cores representative of different HER2 expression levels (null HER2 expression, low HER2 expression and high HER2 expression). |  | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-010 | The Cell Line Control (CLC) does not function as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The functionality demonstrated by the CLC is matrix‑dependent and may differ from that observed in tissue controls | 13 | Verified against supplied workbook row | PR.03 | The product, when used as part of the HER2 BMA CDx assay system, shall be capable of confirming correct execution of the assay when used in accordance with recommended assay parameters. |  | Deviations in the assay reagents, instrument and protocol that alter the HER2 expression status in patient samples should result in a failure of the CLC to meet the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Incorrect interpretation of the cell line controls may result in a run being unnecessarily voided | 24 | Verified against supplied workbook row | PR.03 | The product, when used as part of the HER2 BMA CDx assay system, shall be capable of confirming correct execution of the assay when used in accordance with recommended assay parameters. |  | Deviations in the assay reagents, instrument and protocol that alter the HER2 expression status in patient samples should result in a failure of the CLC to meet the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-009 | The product is not compatible with components of the HER2 BMA Assay System | Product is not compatible with HER2 BMA Assay System and does not stain correctly resulting in a false void run | 12 | Verified against supplied workbook row | PR.04 | The Product must be compatible for use with all components of the HER2 BMA CDx Assay System. | DP61-VEP-006 | The Product, when stained using all components of the HER2 BMA Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained using all components of the HER2 BMA CDx Assay system, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using all components of the HER2 BMA CDx Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-013 | The product IFU and interpretation guide does not explain the scoring system or contain images of expected stained cell lines | The product IFU and interpretation guide is inadequate and CLC failures are not identified | 15 | Verified against supplied workbook row | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | DP61-VEP-002 | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-040 | The product is not able to be evaluated consistently by different pathologists | CLC may not correctly be interpretated by all pathologists | 32 | Verified against supplied workbook row | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | DP61-VEP-002 | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Incorrect interpretation of the cell line controls may result in a run being unnecessarily voided | 24 | Verified against supplied workbook row | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | DP61-VEP-002 | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-013 | The product IFU and interpretation guide does not explain the scoring system or contain images of expected stained cell lines | The product IFU and interpretation guide is inadequate and CLC failures are not identified | 15 | Verified against supplied workbook row | PR.07 | The Product labelling must describe the expected staining patterns and characteristics of each cell cores. | DP61-VEP-002 | HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to confirm they describes the expected staining patterns and characteristics of each cell line core. Sample Size: TBD | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-024 | CLC slide is not evaluated as intended | Incorrect interpretation of the cell line controls may result in a run being unnecessarily voided | 24 | Verified against supplied workbook row | PR.07 | The Product labelling must describe the expected staining patterns and characteristics of each cell cores. | DP61-VEP-002 | HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to confirm they describes the expected staining patterns and characteristics of each cell line core. Sample Size: TBD | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-018 | The product does not maintain acceptable staining performance after transportation to the customer | Product is exposed to extreme conditions during transport and suboptimal staining occurs resulting in a void run | 17 | Verified against supplied workbook row | PR.08 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | DP61-VEP-005 | The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | 100% of stained Product must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide after being subject to transport simulation.The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-001 | The product is not ready to use and requires processing prior to use | Additional processing required to achieve optimal staining and so when used as ready-to-use results in failure to meet acceptance criteria leading to a void run | 8 | Verified against supplied workbook row | PR.09 | The Product must generate Acceptable Staining without having to be subjected to any immediate preparation (e.g., slide baking or incubation) after being removed from the recommended storage conditions (2-8Â°C). | DP61-VEP-006 | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-004 | The product cannot be used directly from recommended storage | The product is loaded onto the BOND straight from storage and fails to meet acceptance criteria leading to a void run | 9 | Verified against supplied workbook row | PR.09 | The Product must generate Acceptable Staining without having to be subjected to any immediate preparation (e.g., slide baking or incubation) after being removed from the recommended storage conditions (2-8Â°C). | DP61-VEP-006 | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | [5 empty] | No risk citation in DP61-PR-001 | PR.10 | The Product must generate Acceptable Staining when used in conjunction with the recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols. | DP61-VEP-006 | The Product, when stained using the HER2 BMA Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-017 | Each cell line does not maintain its HER2 expression level over the shelf life when stored at recommended temperature, and does not have a minimum shelf life of 6 months for initial release | Product is loaded onto BOND within shelf life but suboptimal staining occurs resulting in a void run | 16 | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life when stored in the immediate packaging (minimum of six months shelf life with a minimum two months safety margin). | DP61-VEP-005 | Following a minimum of 14 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. The Product must also retains its ability to correctly detect procedural failures. Sample Size: TBD | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.Following a minimum of 18 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The Product must also retain its ability to correctly detect procedural failures. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-043 | The product (CLC) is used beyond the expiry date | The user uses the product (CLC) beyond the expiry date and suboptimal staining leads to a incorrect void run | 31 | Verified against supplied workbook row | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life when stored in the immediate packaging (minimum of six months shelf life with a minimum two months safety margin). | DP61-VEP-005 | Following a minimum of 14 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. The Product must also retains its ability to correctly detect procedural failures. Sample Size: TBD | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.Following a minimum of 18 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The Product must also retain its ability to correctly detect procedural failures. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-039 | The product is not reproducible and does not have consistent appearance/ morphology between lots | The product cannot be scored due to unacceptable cell line staining and morphology/appearance, resulting in a void run | 31 | Verified against supplied workbook row | PR.12 | The Intermediate Product must be stable for a minimum of six months when stored at room temperature. | DP61-VEP-005 | The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide after being stored at room temperature for at least six months. Sample Size: TBD | Following a minimum of six months storage at room temperature, 100% of The Intermediate Product (individual agarose cell line cores of BT474-M3, RT4 and PANC-1) continues to produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1 must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide after being stored at room temperature for at least six months. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-021 | The product does not maintain HER2 expression levels when different lots of the CLC are used | Product does not maintain HER2 expression levels across different lots of the CLC resulting in a void run | 18 | Verified against supplied workbook row | PR.13 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | DP61-VEP-003 | Three lots of the Product, manufactured from separate multiblocks will be stained using a single lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by one trained pathologist, with the lot of the Product as the only variable. Staining results must align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must demonstrate Acceptable Staining and Consistent Staining across six different lots. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.For inter-lot testing, the lower bound of the two-sided 95% confidence interval for the overall percent agreement should be at least 85%. Testing will be performed on 6 lots. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-039 | The product is not reproducible and does not have consistent appearance/ morphology between lots | The product cannot be scored due to unacceptable cell line staining and morphology/appearance, resulting in a void run | 31 | Verified against supplied workbook row | PR.13 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | DP61-VEP-003 | Three lots of the Product, manufactured from separate multiblocks will be stained using a single lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by one trained pathologist, with the lot of the Product as the only variable. Staining results must align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must demonstrate Acceptable Staining and Consistent Staining across six different lots. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.For inter-lot testing, the lower bound of the two-sided 95% confidence interval for the overall percent agreement should be at least 85%. Testing will be performed on 6 lots. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | [5 empty] | No risk citation in DP61-PR-001 | PR.14 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining throughout the same lots of The Product (intra-lot repeatability). | DP61-VEP-003 | Sections taken from the beginning, middle, and end of a single Cell Line Control multiblock will be stained using one lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. A trained pathologist will evaluate the slides, with the section position within the multiblock as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs For intra-lot testing, taken from within a single lot must generate Acceptable Staining and Consistent Staining. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate the overall percent agreement should be at least 85%. Only sections within a single lot will be tested. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-022 | The product does not maintain HER2 expression levels when stained on multiple BOND-III instruments over multiple days | Product does not maintain HER2 expression levels when different BOND-III machines are used and on different days resulting in a void run | 21 | Verified against supplied workbook row | PR.15 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained on five non-consecutive days (spanning a minimum period of 20 days). A single trained pathologist will score the slides, with the staining day being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained over 5 non-consecutive days (spanning a minimum period of 20 days), CLCs must generate Acceptable Staining and Consistent Staining. For inter-day testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate for the overall percent agreement should should be at least 85%. Testing will occur over 5 non-consecutive days spanning a minimum period of 20 days. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-022 | The product does not maintain HER2 expression levels when stained on multiple BOND-III instruments over multiple days | Product does not maintain HER2 expression levels when different BOND-III machines are used and on different days resulting in a void run | 21 | Verified against supplied workbook row | PR.16 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained using three different BOND-III fully automated systems. A single trained pathologist will score the slides, with the BOND-III instrument being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across three different BOND-IIIs. For inter-instrument testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed across three BOND-IIIs. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | [5 empty] | No risk citation in DP61-PR-001 | PR.17 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained on different BOND Runs/Slide Staining Assemblies (SSAs) within the same BOND Run Batch. A single trained pathologist will score the slides, with the BOND Run/SSA as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across different SSAs on a BOND-III Instrument run batch and within BOND Runs (SSAs). For inter-run testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Iner-run data will be collected across BOND run batches. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-021.2 | The product does not maintain HER2 expression levels when different lots of the detection kit DS9800 are used | Product does not maintain HER2 expression levels across different lots of the detection kit DS9800 resulting in a void run | 20 | Verified against supplied workbook row | PR.18 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Polymer Refine Detection (DS9800) system. The slides will be scored by a single trained pathologist, with the lot of the DS9800 as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system For inter-detectiontested samples lot repeatability testing, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Three lots will be used during testing. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-021.1 | The product does not maintain HER2 expression levels when different lots of the antibody are used | Product does not maintain HER2 expression levels across different lots of the antibody resulting in a void run | 19 | Verified against supplied workbook row | PR.19 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx are varied (inter-antibody lot repeatability). | DP61-VEP-003 | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by a single trained pathologist, with the lot of the antibody as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. For inter-antibody lot repeatability testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed using 3 lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-047 | The product does not maintain HER2 expression levels across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The product does not maintain HER2 expression levels across the range of temperatures defined in the BOND-III user manual leading to a void run | 39 | Verified against supplied workbook row | PR.20 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. |  | Cell Line Control slides from BOND runs performed at environmental temperatures between 18-26°C must stain as expected - Core A (BT474-M3) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-048 | The product does not maintain HER2 expression levels across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The product does not maintain HER2 expression levels across the range of operating humidity's defined in the BOND user manual leading to a void run | 40 | Verified against supplied workbook row | PR.21 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III humidity specifications. |  | Cell Line Control slides from BOND runs performed at environmental humidities between 30-80% Relative Humidity must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-006 | The FFPE cell cores are not presented in every section in the appropriate orientation | Product is loaded onto BOND-III and stained with HER2 BMA antibody, but fails to meet acceptance criteria as user scores the cell cores incorrectly based on expected cell core orientation shown in the Interpretation Guide | 10 | Verified against supplied workbook row | PR.22 | The placement of the cell line cores within The Product must be consistently orientated relative to each other. | DP61-VEP-006 | HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA Assay Interpretation Guide. Sample Size: TBD | A HER2 scoring SME/qualified pathologist will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product IFU and HER2 BMA CDx Assay Interpretation Guide. 100% of assessed CLC slides must demonstrate correct core orientation. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-028 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by GLP or product use may lead to user cutting themselves | 22 | Verified against supplied workbook row | PR.23 | The Product must have labelling that provides all required hazard and safety information and associated symbols. | DP61-VEP-002 | Users will review the Product slide mailer label, hazard notification slip and IFU and complete a checklist to confirm that the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents) is present. Sample Size: TBD | Users will review the Product slide mailer label, IFU and product specifications and complete a checklist to confirm that the required safety information is present. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-028 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by GLP or product use may lead to user cutting themselves | 22 | Verified against supplied workbook row | PR.24 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | DP61-VEP-002 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. Sample Size: TBD | Laboratory users will confirm that the hazards detailed in the IFU and product specifications can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP). | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-032 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | The outer and immediate packaging is damaged by storage at recommended temperature resulting in damage to the product meaning it cannot be used | 26 | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-032.5 | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged or lost during shipping resulting in damage to the product meaning it cannot be used | 27 | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-033 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | The outer and immediate packaging is damaged by storage at recommended temperature resulting in damage to the product which poses a cutting risk to the user | 28 | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-033.5 | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged during shipping resulting in damage which poses a cutting risk to the user | 29 | Verified against supplied workbook row | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | DP61-VEP-002 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-032 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | The outer and immediate packaging is damaged by storage at recommended temperature resulting in damage to the product meaning it cannot be used | 26 | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-032.5 | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged or lost during shipping resulting in damage to the product meaning it cannot be used | 27 | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-033 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | The outer and immediate packaging is damaged by storage at recommended temperature resulting in damage to the product which poses a cutting risk to the user | 28 | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-033.5 | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged during shipping resulting in damage which poses a cutting risk to the user | 29 | Verified against supplied workbook row | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | DP61-VEP-002 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-032.5 | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged or lost during shipping resulting in damage to the product meaning it cannot be used | 27 | Verified against supplied workbook row | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | DP61-VEP-002 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-050 | The product container (slide mailer) does not remain intact during shipment to customer | Cell line cores may become contaminated or damaged and therefore cannot be used | 44 | Verified against supplied workbook row | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | DP61-VEP-002 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-051 | The product container (slide mailer) is not safe to handle | The user could be harmed by sharp edges of container | 45 | Verified against supplied workbook row | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | DP61-VEP-002 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-031 | The product and immediate container labels containing product, lot number and expiry date are not legible | The product and immediate container labels are not legible and the product is ran out of specification, resulting in a void run | 25 | Verified against supplied workbook row | PR.28 | The on-slide printed text must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-041 | The product labels do not have font and symbols that are human legible | The user cannot read the labelling font or the symbols on the product labels so cannot identify the product | 33 | Verified against supplied workbook row | PR.28 | The on-slide printed text must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-043 | The product labels do not have font and symbols that is legible to laboratory users | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | 35 | Verified against supplied workbook row | PR.29 | The Immediate Packaging label must be capable of withstanding shipping and storage within the recommended temperature range and normal use without becoming detached. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remain fully adhered and undamaged. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remained fully adhered and undamaged. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-044 | The product container labels do not remain adhered during the entire life of the product | During the life of the product the labels become un-adhered to the product after exposure to reagents, rendering the user unable to identify the product/relevant safety labelling | 36 | Verified against supplied workbook row | PR.29 | The Immediate Packaging label must be capable of withstanding shipping and storage within the recommended temperature range and normal use without becoming detached. | DP61-VEP-002 | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remain fully adhered and undamaged. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remained fully adhered and undamaged. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-041 | The product labels do not have font and symbols that are human legible | The user cannot read the labelling font or the symbols on the product labels so cannot identify the product | 33 | Verified against supplied workbook row | PR.30 | The symbols and text included on The Product labelling must be human legible. | DP61-VEP-002 | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | [5 empty] | No risk citation in DP61-PR-001 | PR.31 | The Product shall include compliance with UDI specifications. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 23 | Verified against supplied workbook row | PR.32 | The Product labelling must be created to be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 23 | Verified against supplied workbook row | PR.33 | The Product labelling must be created to be compliant with the relevant requirements from IVDR EU/2017/746. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 23 | Verified against supplied workbook row | PR.34 | The Product labelling must be created to be compliant with the relevant Canadian IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-029 | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | 23 | Verified against supplied workbook row | PR.35 | The Product labelling must be created to comply with the relevant UK general IVD as per the UK MDR 2002 labelling regulations. | DP61-VEP-002 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-049 | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps leading to a void run | 41 | Verified against supplied workbook row | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | DP52/DP61-VEP-001 (Environmental) | Cell Line Control slides from BOND runs performed at altitudes up to and including 1600 meters above sea level must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBC | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-004 | The product cannot be used directly from recommended storage | The product is loaded onto the BOND straight from storage and fails to meet acceptance criteria leading to a void run | 9 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-006 | The FFPE cell cores are not presented in every section in the appropriate orientation | Product is loaded onto BOND-III and stained with HER2 BMA antibody, but fails to meet acceptance criteria as user scores the cell cores incorrectly based on expected cell core orientation shown in the Interpretation Guide | 10 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-007 | The product does not contain three cell line cores which demonstrate a staining range (high expression, low expression and negative) | Product does not contain cores including a range of HER2 expression so is unable to act as a procedural control to detect a failed run | 11 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-010 | The Cell Line Control (CLC) does not function as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The functionality demonstrated by the CLC is matrix‑dependent and may differ from that observed in tissue controls | 13 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-011 | The Cell Line Control (CLC) functions as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The product is stained on the BOND-III, but is not compatible with general laboratory clearing and mounting procedures which results in difficulty in interpretation and a false void run | 14 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-013 | The product IFU and interpretation guide does not explain the scoring system or contain images of expected stained cell lines | The product IFU and interpretation guide is inadequate and CLC failures are not identified | 15 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-028 | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by GLP or product use may lead to user cutting themselves | 22 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-033 | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | The outer and immediate packaging is damaged by storage at recommended temperature resulting in damage to the product which poses a cutting risk to the user | 28 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-033.5 | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged during shipping resulting in damage which poses a cutting risk to the user | 29 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-049 | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps leading to a void run | 41 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-049.01 | The product does not maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude | The product does not maintain HER2 expression levels when ran with a user valdiated modified HIER protocol and fails a run which should pass | 42 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-049.1 | The product does not maintain HER2 expression levels and morphology when used above 1600m | The product does maintain HER2 expression levels or morphology above 1600m leading to a valid run | 43 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-050 | The product container (slide mailer) does not remain intact during shipment to customer | Cell line cores may become contaminated or damaged and therefore cannot be used | 44 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-052 | The product is not ran with the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 46 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-011.5 | CLC preparation is not carried out appropriately | User bakes the CLC resulting in sub-optimal staining | 14 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-024.5 | CLC slide is not evaluated as intended | Incorrect interpretation of the CLC may result in a run being mistakenly passed | 25 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-047 | User does not dispose of glass sides appropriately | Glass slides not properly disposed and may cause injury | 34 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-048 | The BOND run is not set up correctly | Covertiles misplaced on the slide leading to sub-optimal staining | 35 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-049 | The BOND run is not set up correctly | Covertiles are not placed on top of the slides leading to sub-optimal staining | 36 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-050 | The BOND run is not set up correctly | Unclean covertiles are used for staining leading to background staining and/or sub-optimal staining | 37 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-052 | The BOND run is not set up correctly | Insufficient BOND ancillary reagents loaded onto BOND leading to BOND run error or delay | 38 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-053 | The user coverslips using an non-appropriate mountant | User coverslips using non-recommended mountant, leading to sub-optimal visualization of staining | 39 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-058 | The BOND-III instrument is not maintained as per IFU and BOND User Manual | The BOND-III instrument is not cleaned and mould/contaminants in tubing lead to sub-optimal staining. | 41 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-066 | An alternative detection kit to that recommended is used with the product to perform staining | Slides stained using alternative detection kit to the recommended DS9800 leading to sub-optimal staining | 43 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-068 | The product is not used for automated staining on the BOND-III staining platform | The product is used on an alternative automated staining platform to the BOND-III leading to sub-optimal staining | 44 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-070 | The product is ran on the BOND-III using a non recommended BOND software version | The BOND-III is on a non-recommended software version leading to sub-optimal staining | 45 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-072 | Assay is not run alongside CLCs | CLCs are not included in run and sub-optimal staining cannot be detected CLCs are not included in run and certain changes in procedural staining performance may not be identified | 46 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-075 | Product is not run alongside recommended tissue controls | Tissue controls are not included in run and sub-optimal staining cannot be detected | 47 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-078 | Product is not run with recommended staining protocol | User selects a non-recommended staining protocol leading to sub-optimal staining | 48 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-080 | Product is not run with recommended negative reagent control | Negative reagent controls are not included in run and sub-optimal staining cannot be detected | 49 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-083 | The product is not stored within the recommended temperature range | User does not store product at recommended temperature range (2-8°C). Following storage the product is used on the BOND-III and results in sub-optimal staining | 50 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-086 | Slides are not dehydrated and cleared as per standard IHC techniques | User does not appropriately dehydrate and clear the sections resulting in bubbles/debris on slides that interfere with slide interpretation | 51 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-089 | Product is not run with recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 52 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-091 | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | User leaves product on the BOND-III for extended periods of time between use, leading to antibody degradation and evaporation | 53 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-096 | The BOND run completes successfully and the user does not unload the SSA in a timely manner | The user does not unload the SSA, the slides is continually hydrated and staining becomes sub-optimal over time | 56 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-097 | User does not use the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | 57 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-euRMW-020 | euRMW-risk-098 | User runs assay at altitudes greater than 1600m | Staining in tissue samples differs to the staining that would have resulted if ran at altitudes below 1600m Epitope retrieval failure leads to sub optimal staining or patchy staining | 58 | Verified against supplied workbook row | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | DP61-VEP-002 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-011 | The Cell Line Control (CLC) functions as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The product is stained on the BOND-III, but is not compatible with general laboratory clearing and mounting procedures which results in difficulty in interpretation and a false void run | 14 | Verified against supplied workbook row | PR.38 | The Product must be compatible with reagent grade alcohol, xylene and toluene based mounting media. |  | All Cell Line Controls must maintain acceptable staining and cell morphology when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA Assay System. All Cell Line Controls must maintain acceptable staining and cell morphology when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA Assay System. Sample Size: TBD |  | [NEEDS INPUT] No supplied protocol coverage row | PR-to-risk: DP61-PR-001 Rev 05; PR-to-protocol: supplied protocol
DP61 | LBN-dRMW-017 | dRMW-risk-049.01 | The product does not maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude |  | 42 | Verified exact ID exists | N/A | The product shall maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude | DP52/DP61-VEP-001 (Environmental) | N/A - risk-characterisation-only protocol row | 100% of sections stained at 1600m using the alternative HIER protocols (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins) should demonstrate no evidence of altitude staining artifact impacting slide interpretation or altered HER2 status as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | N/A - no PR reference in protocol | Risk-characterisation-only row; protocol states PR reference N/A
DP61 | LBN-dRMW-017 | dRMW-risk-049.1 | The product does not maintain HER2 expression levels and morphology when used above 1600m |  | 43 | Verified exact ID exists | N/A | The product shall maintain HER2 expression levels and morphology when used above 1600m | DP52/DP61-VEP-001 (Environmental) | N/A - risk-characterisation-only protocol row | 100% of sections stained at 1800m should demonstrate no evidence of altitude staining artifact impacting slide interpretation or altered HER2 status as defined in Section 5.3.1. The HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. If 100% of sections do not stain satisfactorily, then either the positive/negative tissue controls or CLCs must fail. | N/A - no PR reference in protocol | Risk-characterisation-only row; protocol states PR reference N/A

## Copy to Protocols
ProductFamily | ProtocolID | ProductRequirementID | ProductRequirementText | OverarchingAcceptanceCriteria_CopyReady | CurrentProtocolAcceptanceCriteria | CopyStatus | OverarchingSource | ProtocolSource | ProtocolFile
DP52 | DP52.DP61-VAP-002 | PR.58 | The Product, when used within the HER2 BMA CDx Assay System, is expected to show greater agreement with reflex testing results, helping enable a more definitive determination of HER2 status based on the initial IHC test | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | MATCHES OVERARCHING | DP52-VEP-001, Page 77 | Table 5, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52 | DP52-VEP-003 | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 15-17 | Table 12, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 19-21 | Table 12, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 21-23 | Table 12, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.18 | The Product labelling will detail known limitations of formalin fixed, paraffin-embedded (FFPE) cut section stability. | Product Instructions for Use (IFU) document specifies any known limitations for cut tissue sections during the stability period as determined by verification testing. | Product Instructions For Use (IFU) document specifies the recommended storage conditions for cut tissue sections during the stability period (temperature, humidity, protection from light). | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 30-32 | Table 12, row 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | MATCHES OVERARCHING | DP52-VEP-001, Page 34-36 | Table 12, row 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 55-57 | Table 12, row 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | MATCHES OVERARCHING | DP52-VEP-001, Page 60-63 | Table 12, row 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.44 | The symbols and text included on The Product labelling must be human legible. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | DP52-VEP-001, Page 63 | Table 12, row 9 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | MATCHES OVERARCHING | DP52-VEP-001, Page 64 | Table 12, row 10 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.52 | The Product shall include compliance with UDI specifications. | The labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | The final labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 72-73 | Table 12, row 11 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.53 | The Product labelling must be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | MATCHES OVERARCHING | DP52-VEP-001, Page 73 | Table 12, row 12 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.54 | The Product labelling must be compliant with the relevant requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | MATCHES OVERARCHING | DP52-VEP-001, Page 73-75 | Table 12, row 13 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.55 | The Product labelling must be compliant with the requirements of a Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | MATCHES OVERARCHING | DP52-VEP-001, Page 75 | Table 12, row 14 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.56 | The Product labelling must be compliant with the relevant UK general IVD per the UK MDR 2002 labelling regulations. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | MATCHES OVERARCHING | DP52-VEP-001, Page 76 | Table 12, row 15 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | MATCHES OVERARCHING | DP52-VEP-001, Page 78-80 | Table 12, row 16 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | MATCHES OVERARCHING | DP52-VEP-001, Page 81-83 | Table 12, row 17 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | MATCHES OVERARCHING | DP52-VEP-001, Page 83-86 | Table 12, row 18 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-005 | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | MATCHES OVERARCHING | DP52-VEP-001, Page 14-15 | Table 10, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-005 PREVALENCE Core Team Review.docx
DP52 | DP52-VEP-007 | PR.25 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of HER2 BMA Antibody CDx on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. This acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%.Point estimate of ≥90% when the lower bound of the 95% Confident Interval is ≥85%. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 38-40 | Table 16, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.26 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 40-42 | Table 16, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.27 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different BOND-III systems. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 42-43 | Table 16, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.28 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different SSAs on a BOND-III Instrument run batch. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 44-45 | Table 16, row 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.29 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining within BOND Runs (intra-run repeatability). | For each of the conditions listed below, a minimum of 90% of tissue cases must demonstrate Acceptable Staining and Consistent Staining across the same SSA on a BOND-III Instrument. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 46-47 | Table 16, row 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.30 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 47-48 | Table 16, row 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-008 | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | DP52-VEP-001, Page 18 | Table 12, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-008 | PR.41 | The Product must generate Acceptable Staining when there is a delay of up to seven days between placing The Product on the BOND-III and commencement of staining. | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | MATCHES OVERARCHING | DP52-VEP-001, Page 59 | Table 12, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-008 | PR.42 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining when there is a delay of up to 24 hours between completion of staining and removal. | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | MATCHES OVERARCHING | DP52-VEP-001, Page 59-60 | Table 12, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-008 | PR.64 | The Product must be compatible with reagent‑grade alcohol, xylene, and xylene or toluene‑based mounting media. | All tissue cases must maintain acceptable staining when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA CDx Assay System, all slides must maintain acceptable staining when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA CDx Assay System. | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 86-87 | Table 12, row 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-010 | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | MATCHES OVERARCHING | DP52-VEP-001, Page 18-19 | Table 12, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-010 Sample Characterisation Ver Protocol.docx
DP52 | DP52-VEP-011 | PR.01 | The Product must be specific for the HER2 protein. | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 12-13 | Table 12, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52 | DP52-VEP-012 | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | DP52-VEP-001, Page 28-29 | Table 7, row 2 | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52 | DP52-VEP-012 | PR.10 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim. | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 23-24 | Table 7, row 3 | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52 | DP52-VEP-012 | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | MATCHES OVERARCHING | DP52-VEP-001, Page 24-25 | Table 7, row 4 | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52 | DP52-VEP-013 | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 14 | Table 11, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | MATCHES OVERARCHING | DP52-VEP-001, Page 25-26 | Table 11, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | MATCHES OVERARCHING | DP52-VEP-001, Page 26 | Table 11, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | MATCHES OVERARCHING | DP52-VEP-001, Page 32-33 | Table 11, row 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 33 | Table 11, row 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 33-34 | Table 11, row 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.39 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance.The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. Stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 58 | Table 11, row 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.40 | The Product can be run within the same BOND Run Batch as other BOND Detection systems and staining protocol combinations. | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | MATCHES OVERARCHING | DP52-VEP-001, Page 58-59 | Table 11, row 9 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.45 | The Product when expired, will not allow a staining run to commence. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument.Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 64 | Table 11, row 10 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 77-78 | Table 11, row 11 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP61 | DP61-VEP-002 | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 11; Table 5, row 7 | Table 9, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.07 | The Product labelling must describe the expected staining patterns and characteristics of each cell cores. | HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to confirm they describes the expected staining patterns and characteristics of each cell line core. Sample Size: TBD | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 12; Table 5, row 8 | Table 9, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.23 | The Product must have labelling that provides all required hazard and safety information and associated symbols. | Users will review the Product slide mailer label, hazard notification slip and IFU and complete a checklist to confirm that the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents) is present. Sample Size: TBD | Users will review the Product slide mailer label, IFU and product specifications and complete a checklist to confirm that the required safety information is present. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 17; Table 5, row 24 | Table 9, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.24 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. Sample Size: TBD | Laboratory users will confirm that the hazards detailed in the IFU and product specifications can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP). | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 17; Table 5, row 25 | Table 9, row 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 17; Table 5, row 26 | Table 9, row 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 18; Table 5, row 27 | Table 9, row 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 18; Table 5, row 28 | Table 9, row 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.28 | The on-slide printed text must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired. | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 18; Table 5, row 29 | Table 9, row 9 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.29 | The Immediate Packaging label must be capable of withstanding shipping and storage within the recommended temperature range and normal use without becoming detached. | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remain fully adhered and undamaged. Sample Size: TBD | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remained fully adhered and undamaged. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 18; Table 5, row 30 | Table 9, row 10 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.30 | The symbols and text included on The Product labelling must be human legible. | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. Sample Size: TBD | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 19; Table 5, row 31 | Table 9, row 11 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.31 | The Product shall include compliance with UDI specifications. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 19; Table 5, row 32 | Table 9, row 12 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.32 | The Product labelling must be created to be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 19; Table 5, row 33 | Table 9, row 13 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.33 | The Product labelling must be created to be compliant with the relevant requirements from IVDR EU/2017/746. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 20; Table 5, row 34 | Table 9, row 14 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.34 | The Product labelling must be created to be compliant with the relevant Canadian IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 20; Table 5, row 35 | Table 9, row 15 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.35 | The Product labelling must be created to comply with the relevant UK general IVD as per the UK MDR 2002 labelling regulations. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. Sample Size: TBD | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 20; Table 5, row 36 | Table 9, row 16 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 21; Table 5, row 38 | Table 9, row 17 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-003 | PR.13 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | Three lots of the Product, manufactured from separate multiblocks will be stained using a single lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by one trained pathologist, with the lot of the Product as the only variable. Staining results must align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must demonstrate Acceptable Staining and Consistent Staining across six different lots. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.For inter-lot testing, the lower bound of the two-sided 95% confidence interval for the overall percent agreement should be at least 85%. Testing will be performed on 6 lots. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 14; Table 5, row 14 | Table 10, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.14 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining throughout the same lots of The Product (intra-lot repeatability). | Sections taken from the beginning, middle, and end of a single Cell Line Control multiblock will be stained using one lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. A trained pathologist will evaluate the slides, with the section position within the multiblock as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs For intra-lot testing, taken from within a single lot must generate Acceptable Staining and Consistent Staining. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate the overall percent agreement should be at least 85%. Only sections within a single lot will be tested. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 14; Table 5, row 15 | Table 10, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.15 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | Cell Line Control slides from the same lot will be stained on five non-consecutive days (spanning a minimum period of 20 days). A single trained pathologist will score the slides, with the staining day being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained over 5 non-consecutive days (spanning a minimum period of 20 days), CLCs must generate Acceptable Staining and Consistent Staining. For inter-day testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate for the overall percent agreement should should be at least 85%. Testing will occur over 5 non-consecutive days spanning a minimum period of 20 days. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 14; Table 5, row 16 | Table 10, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.16 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | Cell Line Control slides from the same lot will be stained using three different BOND-III fully automated systems. A single trained pathologist will score the slides, with the BOND-III instrument being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across three different BOND-IIIs. For inter-instrument testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed across three BOND-IIIs. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 15; Table 5, row 17 | Table 10, row 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.17 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | Cell Line Control slides from the same lot will be stained on different BOND Runs/Slide Staining Assemblies (SSAs) within the same BOND Run Batch. A single trained pathologist will score the slides, with the BOND Run/SSA as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across different SSAs on a BOND-III Instrument run batch and within BOND Runs (SSAs). For inter-run testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Iner-run data will be collected across BOND run batches. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 15; Table 5, row 18 | Table 10, row 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.18 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Polymer Refine Detection (DS9800) system. The slides will be scored by a single trained pathologist, with the lot of the DS9800 as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system For inter-detectiontested samples lot repeatability testing, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Three lots will be used during testing. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 15; Table 5, row 19 | Table 10, row 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.19 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx are varied (inter-antibody lot repeatability). | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by a single trained pathologist, with the lot of the antibody as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. For inter-antibody lot repeatability testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed using 3 lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 16; Table 5, row 20 | Table 10, row 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-005 | PR.08 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | 100% of stained Product must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide after being subject to transport simulation.The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 12; Table 5, row 9 | Table 12, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61 | DP61-VEP-005 | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life when stored in the immediate packaging (minimum of six months shelf life with a minimum two months safety margin). | Following a minimum of 14 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. The Product must also retains its ability to correctly detect procedural failures. Sample Size: TBD | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.Following a minimum of 18 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The Product must also retain its ability to correctly detect procedural failures. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 13; Table 5, row 12 | Table 12, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61 | DP61-VEP-005 | PR.12 | The Intermediate Product must be stable for a minimum of six months when stored at room temperature. | The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide after being stored at room temperature for at least six months. Sample Size: TBD | Following a minimum of six months storage at room temperature, 100% of The Intermediate Product (individual agarose cell line cores of BT474-M3, RT4 and PANC-1) continues to produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1 must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide after being stored at room temperature for at least six months. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 13; Table 5, row 13 | Table 12, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61 | DP61-VEP-006 | PR.04 | The Product must be compatible for use with all components of the HER2 BMA CDx Assay System. | The Product, when stained using all components of the HER2 BMA Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained using all components of the HER2 BMA CDx Assay system, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using all components of the HER2 BMA CDx Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 11; Table 5, row 5 | Table 10, row 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.05 | The Product must be compatible for use on the BOND-III fully automated system. |  | The Product, when loaded onto the BOND-III instrument, must not trigger any warning or error messages upon initiation of a staining run. The Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | [NEEDS INPUT] Requirement not found in supplied overarching plan |  | Table 10, row 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.09 | The Product must generate Acceptable Staining without having to be subjected to any immediate preparation (e.g., slide baking or incubation) after being removed from the recommended storage conditions (2-8Â°C). | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 12; Table 5, row 10 | Table 10, row 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.10 | The Product must generate Acceptable Staining when used in conjunction with the recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols. | The Product, when stained using the HER2 BMA Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | When stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 13; Table 5, row 11 | Table 10, row 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.22 | The placement of the cell line cores within The Product must be consistently orientated relative to each other. | HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA Assay Interpretation Guide. Sample Size: TBD | A HER2 scoring SME/qualified pathologist will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product IFU and HER2 BMA CDx Assay Interpretation Guide. 100% of assessed CLC slides must demonstrate correct core orientation. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 16; Table 5, row 23 | Table 10, row 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP52 | DP52/DP61-VEP-001 (Environmental) | PR.35 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96 °C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC | DP52-VEP-001, Page 54 | Table Table 15, row Page 27 | C:\Users\emma.goodrick\Downloads\DP5261-VEP-001_-_Environmental_Verification_Protocol_for_the_HER2_BMA_CDx_Assay Signed 28-July-2026.pdf
DP61 | DP52/DP61-VEP-001 (Environmental) | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | Cell Line Control slides from BOND runs performed at altitudes up to and including 1600 meters above sea level must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBC | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | COPY/RECONCILE OVERARCHING AC | DP61-PR-001 / DP61-VEP-001, PR page 21; Table 5, row 37 | Table Table 15, row Page 27 | C:\Users\emma.goodrick\Downloads\DP5261-VEP-001_-_Environmental_Verification_Protocol_for_the_HER2_BMA_CDx_Assay Signed 28-July-2026.pdf
DP52 | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | DP52-VEP-001, Page 27 | Table RP101-IUVP-001 Section 10; RP101-IUVR-001 Section 5, row Page Plan page 16; report page 11 | T:\R&D\PROJECTS\NPD Overarching Documents\In-Use Verification\RP101-IUVP-001. NPD Overarching In-Use Verification plan.pdf
DP52 | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | MATCHES OVERARCHING | DP52-VEP-001, Page 27-28 | Table RP101-IUVP-001 Section 10; RP101-IUVR-001 Section 5, row Page Plan page 16; report page 11 | T:\R&D\PROJECTS\NPD Overarching Documents\In-Use Verification\RP101-IUVP-001. NPD Overarching In-Use Verification plan.pdf

## Issues
Severity | Category | ProductFamily | DocumentOrRecord | Finding | RecommendedAction
Source conflict | Overarching PR absent from signed PR | DP61 | DP61-VEP-001 PR.01 | [NEEDS INPUT] The overarching plan contains PR.01, but signed DP61-PR-001 Rev 05 does not. | Reconcile the overarching plan with signed DP61-PR-001 Rev 05 under document control.
Source conflict | Overarching PR absent from signed PR | DP61 | DP61-VEP-001 PR.05 | [NEEDS INPUT] The overarching plan contains PR.05, but signed DP61-PR-001 Rev 05 does not. | Reconcile the overarching plan with signed DP61-PR-001 Rev 05 under document control.
Source conflict | Overarching PR absent from signed PR | DP61 | DP61-VEP-001 PR.39 | [NEEDS INPUT] The overarching plan contains PR.39, but signed DP61-PR-001 Rev 05 does not. | Reconcile the overarching plan with signed DP61-PR-001 Rev 05 under document control.
Source conflict | Overarching PR absent from signed PR | DP61 | DP61-VEP-001 PR.40 | [NEEDS INPUT] The overarching plan contains PR.40, but signed DP61-PR-001 Rev 05 does not. | Reconcile the overarching plan with signed DP61-PR-001 Rev 05 under document control.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 57 | [NEEDS INPUT] Source identifier says dMRW; canonicalized to dRMW for linkage: raw='dMRW-risk-086', canonical='dRMW-risk-086' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 58 | Source cell contains identifier variants; canonical valid identifier selected: raw='dMRW-risk-087 dRMW-risk-087', canonical='dRMW-risk-087' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 59 | Source cell contains identifier variants; canonical valid identifier selected: raw='dMRW-risk-088 dRMW-risk-088', canonical='dRMW-risk-088' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 60 | Source cell contains identifier variants; canonical valid identifier selected: raw='dMRW-risk-089 dRMW-risk-089', canonical='dRMW-risk-089' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 61 | Source cell contains identifier variants; canonical valid identifier selected: raw='dMRW-risk-090 dRMW-risk-090', canonical='dRMW-risk-090' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 62 | Source cell contains identifier variants; canonical valid identifier selected: raw='dMRW-risk-091 dRMW-risk-091', canonical='dRMW-risk-091' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 63 | Source cell contains identifier variants; canonical valid identifier selected: raw='dMRW-risk-092 dRMW-risk-092', canonical='dRMW-risk-092' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 64 | Source cell contains identifier variants; canonical valid identifier selected: raw='dMRW-risk-093 dRMW-risk-093', canonical='dRMW-risk-093' | Confirm and correct the controlled risk-line identifier at source.
Source identifier | Risk ID variant | DP52/System | LBN-dRMW-016, Excel row 65 | Source cell contains identifier variants; canonical valid identifier selected: raw='dRMW-risk-094 dRMW-risk-094', canonical='dRMW-risk-094' | Confirm and correct the controlled risk-line identifier at source.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 12, PR.02 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 13, PR.05 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 14, PR.06 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 14, PR.07 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 14, PR.08 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 18, PR.21 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 20, PR.26 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 32, PR.59 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 32, PR.59 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 33, PR.62 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | PR risk citation not found | DP52 | DP52-PR-001 Rev 06 page 33, PR.62 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the PR citation with the current risk workbook.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 3, PR.07 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 4, PR.08 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 7, PR.38 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-095 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 10, PR.46 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-030 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 10, PR.46 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-095 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 17, PR.62 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 17, PR.62 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-003, table 12, row 17, PR.62 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-095 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-005, table 10, row 2, PR.03 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-008, table 12, row 2, PR.05 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-010, table 12, row 2, PR.06 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 2, PR.02 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 2, PR.02 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-025 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 6, PR.20 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-034 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 7, PR.21 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 7, PR.21 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-034 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 7, PR.21 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-088 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 11, PR.59 | [NEEDS INPUT] LBN-dRMW-016 dRMW-risk-030.1 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 11, PR.59 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-001 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP52 | DP52-VEP-013, table 11, row 11, PR.59 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-095 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-002, table 9, row 6, PR.25 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-035 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-002, table 9, row 6, PR.25 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-037 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-002, table 9, row 7, PR.26 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-035 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-002, table 9, row 8, PR.27 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-027 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-002, table 9, row 8, PR.27 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-035 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-002, table 9, row 17, PR.37 | [NEEDS INPUT] LBN-euRMW-020 euRMW-risk-095 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-006, table 10, row 4, PR.09 | [NEEDS INPUT] LBN-dRMW-017 dRMW-risk-019 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-006, table 10, row 4, PR.09 | [NEEDS INPUT] LBN-dRMW-017 dRMW-risk-026 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Traceability gap | Protocol risk citation not found | DP61 | DP61-VEP-006, table 10, row 5, PR.10 | [NEEDS INPUT] LBN-dRMW-017 dRMW-risk-026 was cited but not found in the supplied workbook. | Reconcile the protocol citation with the current risk workbook and overarching/PR source.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-003, PR.04 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-003, PR.07 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-003, PR.08 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-003, PR.18 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-003, PR.38 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-003, PR.52 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-007, PR.25 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-007, PR.26 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-007, PR.27 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-007, PR.28 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-007, PR.29 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-007, PR.30 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-008, PR.64 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-011, PR.01 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-012, PR.10 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-013, PR.02 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-013, PR.20 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-013, PR.21 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-013, PR.39 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-013, PR.45 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52-VEP-013, PR.59 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.06 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.07 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.23 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.24 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.25 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.26 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.27 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.28 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.29 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.30 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.31 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.32 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.33 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.34 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.35 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-002, PR.37 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-003, PR.13 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-003, PR.14 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-003, PR.15 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-003, PR.16 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-003, PR.17 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-003, PR.18 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-003, PR.19 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-005, PR.08 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-005, PR.11 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-005, PR.12 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-006, PR.04 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-006, PR.05 | [NEEDS INPUT] Requirement not found in supplied overarching plan | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-006, PR.09 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-006, PR.10 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP61-VEP-006, PR.22 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP52 | DP52/DP61-VEP-001 (Environmental), PR.35 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.
Acceptance criteria | Protocol copy/reconciliation | DP61 | DP52/DP61-VEP-001 (Environmental), PR.36 | COPY/RECONCILE OVERARCHING AC | Review the copy-ready overarching acceptance criteria and update the protocol under document control.

## Risk Register (hidden)
SourceDocument | SourceRevision | SourceExcelRow | RiskID | RawRiskID | IdentifierStatus | SharedRiskID | Section | Function | Hazard | HazardousSituation | ConditionsOfOccurrence | Harm | Severity | PotentialCauses | CurrentControlMeasures
LBN-dRMW-016 | 7 | 8 | dRMW-risk-001 | dRMW-risk-001 | Exact source identifier | N/A | HER2 Antibody | Antibody detects HER2 expression in breast carcinoma FFPE tissue | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to a false negative but cell line controls meet the acceptance criteria | Normal | False Negative or Delay to diagnosis (more than 10 days) | 4 | Pre-analytical conditions (i.e. tissue fixation) prevent appropriate staining | Regional standard guidelines in place for tissue fixation, e.g. ASCO-CAP, CLSI, ESMO
LBN-dRMW-016 | 7 | 9 | dRMW-risk-002 | dRMW-risk-002 | Exact source identifier | N/A | HER2 Antibody | Antibody detects HER2 expression in breast carcinoma FFPE tissue | Antibody is not specific and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | Normal | False Negative, False Negative (HER2-low), False Negative (HER2-ultralow) | 4 | Inappropriate design of antibody | Clone selected is monoclonal and has been sequenced-verified as specific to HER2
LBN-dRMW-016 | 7 | 10 | dRMW-risk-003 | dRMW-risk-003 | Exact source identifier | N/A | HER2 Antibody | Antibody detects HER2 expression in breast carcinoma FFPE tissue | Antibody is not sensitive and does not detect HER2 in intended tissue | No staining present during scoring leading to false negative but cell line controls meet the acceptance criteria | Normal | False Negative, False Negative (HER2-low), False Negative (HER2-ultralow) | 4 | Antibody not sensitive enough to detect HER2 at low/ultralow expression | Clone selection is aligned with current on-market product
LBN-dRMW-016 | 7 | 11 | dRMW-risk-007 | dRMW-risk-007 | Exact source identifier | N/A | HER2 Antibody | Antibody detects HER2 expression in breast carcinoma FFPE tissue | Antibody is not specific/sensitive and does not detect HER2 in intended tissue | No staining present and the cell line control voids the run | Normal | Delay to diagnosis (less than 10 days) | 2 | Pre-analytical conditions (i.e. tissue fixation) prevent appropriate staining | Standard ASCO-CAP guidelines in place for tissue fixation
LBN-dRMW-016 | 7 | 12 | dRMW-risk-010 | dRMW-risk-010 | Exact source identifier | N/A | HER2 Antibody | Antibody detects HER2 expression in breast carcinoma FFPE tissue | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) | 4 | Pre-analytical conditions (i.e. tissue fixation) prevent appropriate staining | Standard ASCO-CAP guidelines in place for tissue fixation
LBN-dRMW-016 | 7 | 13 | dRMW-risk-011 | dRMW-risk-011 | Exact source identifier | N/A | HER2 Antibody | Antibody detects HER2 expression in breast carcinoma FFPE tissue | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present during scoring leading to false positive but cell line controls meet the acceptance criteria | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) | 4 | Inappropriate design of Antibody for intended tissue | The clone selected is monoclonal and has been sequenced as specific for HER2
LBN-dRMW-016 | 7 | 14 | dRMW-risk-016 | dRMW-risk-016 | Exact source identifier | N/A | HER2 Antibody | Antibody detects HER2 expression in breast carcinoma FFPE tissue | Antibody is not specific and wrongly detects HER2 in intended tissue | Background/ non-specific staining present and cell lines void the run | Normal | Delay to diagnosis (less than 10 days) | 2 | Pre-analytical conditions (i.e. tissue fixation) prevent appropriate staining | Standard ASCO-CAP guidelines in place for tissue fixation/processing
LBN-dRMW-016 | 7 | 15 | dRMW-risk-019 | dRMW-risk-019 | Exact source identifier | N/A | HER2 Antibody | The product shall contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained, and the lab does not have additional stock of antibody units | Normal | Delay to diagnosis (less than 10 days) | 2 | The product is designed with inadequate volume to accommodate reagent | None
LBN-dRMW-016 | 7 | 16 | dRMW-risk-020 | dRMW-risk-020 | Exact source identifier | SR-004 | HER2 Antibody | The product shall contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained, and the lab does not have additional stock of antibody units | Fault | Delay to diagnosis (less than 10 days) | 2 | Reagent volume dispensed by BOND III instrument is too high | BOND tracks and updates reagents volumes after each aspiration, and uses the current volumes for allocation (SRCM-006) Product designed using standard BOND containers and standard fill volumes
LBN-dRMW-016 | 7 | 17 | dRMW-risk-020.5 | dRMW-risk-020.5 | Exact source identifier | SR-016 | HER2 Antibody | The product shall contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Insufficient volume dispensed on slide leading to poor staining quality | Fault | Delay to diagnosis (less than 10 days) | 2 | Reagent volume dispensed by BOND III instrument is too low | BOND tracks and updates reagents volumes after each aspiration, and uses the current volumes for allocation (SRCM-006) BOND will detect and report/flag a failed aspirates or dispense and will be displayed in the run events, including liquid level detection errors (SRCM-026) Product designed using standard BOND containers and standard fill volumes
LBN-dRMW-016 | 7 | 18 | dRMW-risk-021 | dRMW-risk-021 | Exact source identifier | SR-005 | HER2 Antibody | The product shall contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | Laboratory unable to achieve the desired number of patients samples to be stained | Fault | Delay to diagnosis (less than 10 days) | 2 | Error with the liquid level sensor registering the correct volume of reagent (too little/too low reagent) | BOND tracks and updates reagents volumes after each aspiration, and uses the current volumes for allocation (SRCM-036)
LBN-dRMW-016 | 7 | 19 | dRMW-risk-021.5 | dRMW-risk-021.5 | Exact source identifier | SR-005 | HER2 Antibody | The product shall contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) | The product does not contain sufficient reagents to stain a minimum of 46 slides (7mL format) and 200 slides (30mL format) leading to an incorrectly processed slide | Laboratory unable to achieve the desired number of patients samples to be stained | Fault | False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Error with the liquid level sensor registering the correct volume of reagent -- trying to aspirate reagent that is not present (too high) | Product designed using standard BOND container and standard fill volumes, which include a overfill volume to account for LLS inaccuracies (SRCM-036). BOND generates an event if the volume detected is higher than the tracked volume (SRCM-012).
LBN-dRMW-016 | 7 | 20 | dRMW-risk-023 | dRMW-risk-023 | Exact source identifier | SR-022 | HER2 Antibody | The product is only compatible for use on the BOND-III instrument | The product is used on a non-BOND-III instrument | The product is loaded onto a non- BOND-III instrument and accepted | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to be used on non-BOND-III Instrument | None
LBN-dRMW-016 | 7 | 21 | dRMW-risk-024 | dRMW-risk-024 | Exact source identifier | SR-006 | HER2 Antibody | The product is used on BOND software version (software 8.x or greater) | The product is used on a BOND software version software version 8.x or greater) | The product is loaded onto a BOND III instrument and is accepted and ran on a non-validated version of BOND software | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | BDD not correctly updated to include product | None
LBN-dRMW-016 | 7 | 22 | dRMW-risk-025 | dRMW-risk-025 | Exact source identifier | N/A | HER2 Antibody | The product shall be developed to be compatible with existing BOND-III System bulk reagents and ancillaries | The product is not compatible with existing BOND-III system bulk reagents and ancillaries | The product stains inappropriately using current BOND-III bulk reagents and ancillaries | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Product not designed to be compatible with BOND bulk reagents and ancillaries | None
LBN-dRMW-016 | 7 | 23 | dRMW-risk-028 | dRMW-risk-028 | Exact source identifier | N/A | HER2 Antibody | The product Interpretation Guide explains the scoring system along with images of expected HER2 expressing breast carcinoma tissue sections | The product Interpretation Guide does not explain the scoring system or contain images of expected HER2 expressing breast carcinoma tissue sections | The Interpretation Guide is inadequate to facilitate HER2 scoring | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Interpretation Guide is not clear to facilitate HER2 scoring | None
LBN-dRMW-016 | 7 | 24 | dRMW-risk-029 | dRMW-risk-029 | Exact source identifier | N/A | HER2 Antibody | The product shall maintain comparable and acceptable staining performance when different lots of BMA antibody are used | The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Tissue sections are stained sub-optimally when different lots of antibody are used | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to give reproducible staining when different lots of antibody are used | None
LBN-dRMW-016 | 7 | 26 | dRMW-risk-034 | dRMW-risk-034 | Exact source identifier | N/A | HER2 Antibody | The product must be compatible with recommended BOND staining protocols | The product is not compatible with the recommended BOND staining protocols | The product stains sub-optimally using the recommended staining protocol | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to be compatible with recommended staining protocol | BOND run compatibility rules followed to ensure cross-compatibility with other BOND protocols
LBN-dRMW-016 | 7 | 27 | dRMW-risk-039 | dRMW-risk-039 | Exact source identifier | N/A | HER2 Antibody | The product must maintain comparable and acceptable performance when scored by different pathologists | The product does not maintain comparable and acceptable performance when scored by different pathologists | Staining of the tissue is interpretated differently by different pathologists | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed with a scoring approach that gives reproducible results when interpretated by different pathologists | Assay incorporates pertinent ASCO-CAP recommendations
LBN-dRMW-016 | 7 | 28 | dRMW-risk-040 | dRMW-risk-040 | Exact source identifier | N/A | HER2 Antibody | The product must maintain comparable and acceptable performance to the same pathologist over time | The product does not maintain comparable and acceptable performance to the same pathologist over time | Staining of the tissue is interpretated differently by the same pathologist on a different day leading to inconsistent diagnosis | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed with a scoring approach that gives reproducible results when interpretated by the same pathologist over time | Assay incorporates pertinent ASCO-CAP recommendations
LBN-dRMW-016 | 7 | 29 | dRMW-risk-042 | dRMW-risk-042 | Exact source identifier | N/A | HER2 Antibody | The product shall maintain comparable and acceptable staining performance across inter staining runs on the BOND-III instrument | The product does not maintain comparable and acceptable staining performance across both inter staining runs on the BOND-III instrument | The product results in sub-optimal staining when used on different BOND III instruments | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to give reproducible staining when tissue is stained on an alternative BOND III instrument | BOND-III instruments are built and validated to produce scores within 1 point of the average of intensity scores
LBN-dRMW-016 | 7 | 30 | dRMW-risk-043 | dRMW-risk-043 | Exact source identifier | N/A | HER2 Antibody | The product shall maintain comparable and acceptable staining performance across intra staining runs on BOND-III | The product does not maintain comparable and acceptable staining performance across intra staining runs across SSAs on the same BOND-III instrument | The product is loaded onto the BOND and results in sub-optimal staining between different SSAs within the same BOND-III instrument | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to give reproducible staining when tissue is stained on different SSAs within the same BOND | BOND-III instruments are built and validated to produce scores within 1 point of the average of intensity scores
LBN-dRMW-016 | 7 | 31 | dRMW-risk-046 | dRMW-risk-046 | Exact source identifier | N/A | HER2 Antibody | The product must maintain comparable and acceptable performance when used across different days | The product does not maintain comparable and acceptable performance when used across different days | The product is loaded onto the BOND-III and results in sub-optimal staining when used on a different day | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to give reproducible staining when testing is performed on different days | None
LBN-dRMW-016 | 7 | 32 | dRMW-risk-050 | dRMW-risk-050 | Exact source identifier | N/A | HER2 Antibody | The product shall maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | The product does not maintain acceptable staining performance over the life of the reagent on the BOND-III instrument | Staining becomes sub-optimal over the life of the reagent due to dilution of reagents | Normal | False Negative, False Negative (HER2-low), False Negative (HER2-ultralow) | 4 | Aspirating probe is rinsed between reagents and can lead to reagent carry over of wash fluid and dilution of reagents | BOND User manual instructs user to include controls to detect sub-optimal staining (SRCM-025)
LBN-dRMW-016 | 7 | 33 | dRMW-risk-053 | dRMW-risk-053 | Exact source identifier | N/A | HER2 Antibody | The product shall maintain acceptable staining performance over the life of the antibody on the BOND III instrument | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | Staining becomes sub-optimal over the life of the reagent due to evaporation | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) | 4 | Reagents left on BOND in-between use Environmental temperature exceed those of recommended temperatures | None
LBN-dRMW-016 | 7 | 34 | dRMW-risk-056 | dRMW-risk-056 | Exact source identifier | SR-017 | HER2 Antibody | The product shall be able to run on the BOND-III System at the same time as other antibodies and probes | The product is not able to run on the BOND-III System at the same time as other antibodies and probes | The product is loaded onto the BOND-III instrument and is unable to schedule testing alongside other antibodies and probes. Laboratory has to wait until run is complete before they can run the HER2 BMA CDx Assay | Normal | Delay to diagnosis (less than 10 days) | 2 | BOND software cannot schedule the run alongside other expected compatible protocols | BOND run compatibility rules followed to ensure cross-compatibility with other BOND protocols BOND error notifications would occur if scheduling issue
LBN-dRMW-016 | 7 | 35 | dRMW-risk-057 | dRMW-risk-057 | Exact source identifier | SR-018 | HER2 Antibody | The product shall be able to run on the BOND-III System at the same time as other antibodies and probes | The product is not able to run on the BOND-III System at the same time as other antibodies and probes | The product is loaded onto the BOND-III instrument and it completes testing with BOND errors alongside other antibodies and probes | Normal | Delay to diagnosis (less than 10 days) | 2 | Protocol is not designed with adequate extension times per each step | BOND run compatibility rules followed to ensure cross-compatibility with other BOND protocols BOND error notifications would occur if scheduling issue
LBN-dRMW-016 | 7 | 36 | dRMW-risk-058 | dRMW-risk-058 | Exact source identifier | SR-002 | HER2 Antibody | The product is only used within its stated shelf life | The product is used after it is no longer stable | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product becomes unstable over time | BOND software will not allow use of expired reagents (SRCM-001)
LBN-dRMW-016 | 7 | 37 | dRMW-risk-059 | dRMW-risk-059 | Exact source identifier | N/A | HER2 Antibody | The product shall have a minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | The product does not have minimum of stability to allow for production time and a minimum of 6 months stability at the recommended storage temperature | The product does not have the minimum stability | Normal | False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Composition of the product is not suitable | Product diluted using standard diluent which has established stability Product designed using standard BOND containers
LBN-dRMW-016 | 7 | 38 | dRMW-risk-060 | dRMW-risk-060 | Exact source identifier | N/A | HER2 Antibody | The product shall maintain acceptable staining performance after undergoing transportation to the customer | The product does not maintain acceptable staining performance after undergoing transportation to the customer | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Extreme conditions could lead to reagent instability and/or degradation | Product to be shipped following existing shipping procedures
LBN-dRMW-016 | 7 | 39 | dRMW-risk-062 | dRMW-risk-062 | Exact source identifier | N/A | HER2 Antibody | The product shall be able to be used from recommended storage and immediately commence staining operation without a warm-up period | The product cannot be to be used from recommended storage and immediately commence staining operation without a warm-up period | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to produce optimal staining when used directly from recommended storage | None
LBN-dRMW-016 | 7 | 40 | dRMW-risk-066 | dRMW-risk-066 | Exact source identifier | N/A | HER2 Antibody | The product will be ready to use and does not require mixing or shaking before use | The product is not ready to use and requires mixing or shaking before use | The product requires mixing or shaking to produce optimal staining | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The antibody units require mixing or shaking before use | None
LBN-dRMW-016 | 7 | 41 | dRMW-risk-068 | dRMW-risk-068 | Exact source identifier | N/A | Health and Safety | The product shall present no hazards to the users that cannot be adequately controlled by good laboratory practices | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by good laboratory practices | Normal | User harm | 3 | User unaware of hazards | Professional use environment Appropriate PPE is required and routinely used in accordance with standard laboratory procedures
LBN-dRMW-016 | 7 | 42 | dRMW-risk-069 | dRMW-risk-069 | Exact source identifier | N/A | Stability | The product shall have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | The product does not have a minimum in-use stability to stain 16 BOND runs (for 7mL product format) and 23 BOND runs (for 30mL product format) | The product is loaded onto the BOND III instrument for 16 & 23 runs and is not stable for all cycles of use | Normal | False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product is not stable for 16 & 23 cycles of use | Product diluted using standard diluent which has established stability Product designed using standard BOND containers
LBN-dRMW-016 | 7 | 43 | dRMW-risk-070 | dRMW-risk-070 | Exact source identifier | SR-013 | Labelling and Packaging | The product shall include a barcode to allow the BOND System to recognise the identity of the antibody package by means of the on-board scanner | The product does not have a barcode to allow the BOND System to recognise the identity of the antibody package by means of the on-board scanner | The product is loaded onto the BOND-III Instrument and the antibody cannot be recognised by the on-board scanner and cannot automatically be run | Normal | Delay to diagnosis (less than 10 days) | 2 | No barcode included on label design Barcode has not been verified in PRISYM system and so is not recognised by BOND Label design obscures barcode, not readable by BOND | BOND software configured to not allow run to be performed without successful scanning of reagents (SRCM-015) BOND System allows reagents to be identified automatically using the reagent's identification barcode (SRCM-021) Product can be manually registered (SRCM-022) and identified (SRCM-023) into BOND if the barcode is damaged.
LBN-dRMW-016 | 7 | 44 | dRMW-risk-070.5 | dRMW-risk-070.5 | Exact source identifier | SR-012 | Labelling and Packaging | The product shall include a barcode to allow the BOND System to recognise the identity of the reagent package by means of the hand held scanner | The product does not have a barcode to allow the BOND System to recognise the identity of the reagent package by means of the hand held scanner | The product cannot be registered into the BOND inventory by scanning the side barcode | Fault | Delay to diagnosis (less than 10 days) | 2 | No barcode included on label design Barcode has not be verified in PRISYM system and so is not recognised by BOND | BOND System allows reagents to be registered automatically using the reagent's registration barcode (SRCM-021) BOND System allows reagents to be manually registered if the hand-held scanner does not recognise the registration barcode (SRCM-022)
LBN-dRMW-016 | 7 | 45 | dRMW-risk-071 | dRMW-risk-071 | Exact source identifier | N/A | Labelling and Packaging | The product shall have labels on the immediate containers and outer packaging (if required) which will comply with local territory regulations | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | Fault | User harm | 3 | Territories are not considered during label design and development | Regulatory classification considered early in development
LBN-dRMW-016 | 7 | 46 | dRMW-risk-072 | dRMW-risk-072 | Exact source identifier | N/A | Labelling and Packaging | The product labels should have font and symbols that is legible to laboratory users | The product labels do not have font and symbols that is legible to laboratory users | The product is opened and the user cannot read the labelling font or the symbols on the container labels so cannot identify the product | Fault | Delay to diagnosis (less than 10 days) | 2 | The product container labels designed with font and symbols too small to be read | None
LBN-dRMW-016 | 7 | 47 | dRMW-risk-073 | dRMW-risk-073 | Exact source identifier | N/A | Labelling and Packaging | The product labels should have font, symbols and barcode that is legible to laboratory users | The product labels do not have font, symbols and barcode that is legible to laboratory users | The product is opened and the user cannot read the labelling font, symbols or barcode on the container labels so cannot identify the product | Fault | Delay to diagnosis (less than 10 days) | 2 | Font and symbols on the labels is not durable and can be rubbed off during use/manufacture leading to illegible font/symbols | Label material/ink used for all current LBS antibody products
LBN-dRMW-016 | 7 | 48 | dRMW-risk-074 | dRMW-risk-074 | Exact source identifier | N/A | Labelling and Packaging | The product labels should have font and symbols that is legible to laboratory users | The product labels do not have font and symbols that is legible to laboratory users | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | Fault | Delay to diagnosis (less than 10 days) | 2 | The product outer packaging labels designed with font and symbols too small to be read | None
LBN-dRMW-016 | 7 | 49 | dRMW-risk-075 | dRMW-risk-075 | Exact source identifier | N/A | Labelling and Packaging | The product labels should have font and symbols that is legible to laboratory users | The product labels do not have font and symbols that is legible to laboratory users | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | Fault | Delay to diagnosis (less than 10 days) | 2 | Outer packaging labelling font and symbols are not designed to be durable for routine handling by manufacturer/user | None
LBN-dRMW-016 | 7 | 50 | dRMW-risk-076 | dRMW-risk-076 | Exact source identifier | N/A | Labelling and Packaging | The product container labels shall remain adhered during the entire life of the product | The product container labels do not remain adhered during the entire life of the product | During the life of the product the labels become un-adhered to the product after exposure to product reagents or BOND reagents, rendering the user unable to identify the reagent/relevant safety labelling | Fault | User harm | 3 | The labels are not sufficiently adhesive to withstand the normal use of the product throughout its life Reagents from the product may reduce label adhesion | Label material/ink used for all current LBS antibody products
LBN-dRMW-016 | 7 | 51 | dRMW-risk-078 | dRMW-risk-078 | Exact source identifier | N/A | Labelling and Packaging | The product container label text and graphics shall be human readable after the product is exposed to product, BOND or detection kit reagents | The product container label text and graphics are not human readable after the product is exposed to product, BOND or detection kit reagents | The user cannot read the labelling font or the symbols on the container labels after exposure to product reagents or BOND reagents, rendering the user unable to identify relevant safety labelling | Fault | User harm | 3 | The label ink used is not sufficiently durable to withstand exposure to product, BOND or detection kit reagents | Label material/ink used for all current LBS antibody products
LBN-dRMW-016 | 7 | 52 | dRMW-risk-080 | dRMW-risk-080 | Exact source identifier | N/A | Labelling and Packaging | The product container label text and graphics shall be BOND readable after the product is exposed to product, BOND or detection kit reagents | The product container label text and graphics are not BOND readable after the product is exposed to product, BOND or detection kit reagents | The BOND III instrument is unable to read the barcodes of the products and therefore the BOND run cannot schedule | Fault | Delay to diagnosis (less than 10 days delay) | 2 | The label ink used is not sufficiently durable to withstand exposure to product, BOND or detection kit reagents | Label material/ink used for all current LBS antibody products. Users can manually enter the barcode into the inventory
LBN-dRMW-016 | 7 | 53 | dRMW-risk-081 | dRMW-risk-081 | Exact source identifier | N/A | Labelling and Packaging | The product packaging labels shall remain legible after exposure to normal storage and handling conditions | The product packaging labels are not legible after exposure to normal storage and handling conditions | After normal storage and handling the labels are not legible by the user, rendering the user unable to identify the reagent/relevant safety labelling | Normal | User harm | 3 | The label ink used is not sufficiently durable to withstand normal storage and handling | Label material/ink used for all current LBS antibody products
LBN-dRMW-016 | 7 | 54 | dRMW-risk-082 | dRMW-risk-082 | Exact source identifier | N/A | Labelling and Packaging | The product packaging labels shall remain legible after exposure to normal storage and handling conditions | The product packaging labels are not legible after exposure to normal storage and handling conditions | After normal storage and handling the labels are not legible by the BOND system | Normal | Delay to diagnosis (less than 10 days delay) | 2 | The label ink used is not sufficiently durable to withstand normal storage and handling | Label material/ink used for all current LBS antibody products Users can manually enter the barcode into the inventory
LBN-dRMW-016 | 7 | 55 | dRMW-risk-084 | dRMW-risk-084 | Exact source identifier | N/A | Labelling and Packaging | The product shall withstand and not leak, after transportation to the customer | The product does not withstand and/or leaks, after transportation to the customer | The product leaks after transport, exposing the user to the antibody | Normal | User harm | 3 | Packaging of reagents is not robust enough to withstand transportation Insufficient seals on containers Inappropriate handling during transport | Product intermediate packaging is heat sealed
LBN-dRMW-016 | 7 | 56 | dRMW-risk-085 | dRMW-risk-085 | Exact source identifier | N/A | Labelling and Packaging | The product shall withstand and not leak, after transportation to the customer | The product does not withstand and/or leaks, after transportation to the customer | The product reagents leak after transport, reducing the working volume of the product | Normal | Delay to diagnosis (less than 10 days delay) | 2 | Packaging of reagents is not robust enough to withstand transportation Insufficient seals on containers Inappropriate handling during transport | The product is designed using standard BOND containers with intended use for the BOND. Heat seals on containers to prevent leakage BOND LLS measurement on aspiration to determine the volume and alert run event if less than expected volume is present - residual volume determined in container and run continued
LBN-dRMW-016 | 7 | 57 | dRMW-risk-086 | dMRW-risk-086 | [NEEDS INPUT] Source identifier says dMRW; canonicalized to dRMW for linkage | SR-021 | Assay performance | The product shall provide acceptable staining across the range of altitudes defined in the BOND-III user manual, 0-1600m | The product does not provide acceptable staining across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not provide acceptable staining across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps | Normal | False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Customers at higher altitudes work in environments with lower atmospheric pressure, this reduces the boiling point which can affect the effectiveness of the HIER steps for unmasking the epitope of FFPE tissue as these steps are performed at 100°C with the HER2 BMA CDx Assay | None
LBN-dRMW-016 | 7 | 58 | dRMW-risk-087 | dMRW-risk-087 dRMW-risk-087 | Source cell contains identifier variants; canonical valid identifier selected | N/A | HER2 Antibody | The product is compatible with clearing procedures using reagent grade alcohol, xylene, and a xylene/toluene-based mounting media | The product is not compatible with clearing procedures using reagent grade alcohol, xylene, and a xylene/toluene-based mounting media | The product is stained on BOND but is not compatible with laboratory clearing and mounting procedures which results in difficulty in interpretation | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed to be compatible with laboratory clearing and mounting procedures or workflow | None
LBN-dRMW-016 | 7 | 59 | dRMW-risk-088 | dMRW-risk-088 dRMW-risk-088 | Source cell contains identifier variants; canonical valid identifier selected | SR-007 | Assay performance | The HER2 BMA CDx Assay shall provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The HER2 BMA CDx Assay is not designed to work across the range of BOND-III environmental conditions | BOND user manual specifies environmental conditions that the BOND-III has been validated under (SRCM-004)
LBN-dRMW-016 | 7 | 60 | dRMW-risk-089 | dMRW-risk-089 dRMW-risk-089 | Source cell contains identifier variants; canonical valid identifier selected | SR-007 | Assay performance | The HER2 BMA CDx Assay shall provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidities defined in the BOND user manual | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The HER2 BMA CDx Assay is not designed to work across the range of BOND-III environmental conditions | BOND user manual specifies environmental conditions that the BOND-III has been validated under (SRCM-004)
LBN-dRMW-016 | 7 | 61 | dRMW-risk-090 | dMRW-risk-090 dRMW-risk-090 | Source cell contains identifier variants; canonical valid identifier selected | N/A | Labelling and Packaging | Outer packaging (pouch) remains intact during transit to customer | Outer packaging (pouch) does not remain intact during transit to customer | Hazard notification slip may fall out and user won't be aware of hazards and know how to find IFU | Fault | User harm | 3 | Packaging is not robust to withstand standard shipping to customer | The product packaging is designed using standard Leica packaging specifications
LBN-dRMW-016 | 7 | 62 | dRMW-risk-091 | dMRW-risk-091 dRMW-risk-091 | Source cell contains identifier variants; canonical valid identifier selected | N/A | Labelling and Packaging | Outer packaging (pouch) remains intact during transit to customer | Outer packaging (pouch) does not remain intact during transit to customer | Hazard notification slip and products may be exposed to environmental condition affecting legibility of text and labels | Fault | User harm | 3 | Packaging is not robust to withstand standard shipping to customer. | The product packaging is designed using standard Leica packaging specifications
LBN-dRMW-016 | 7 | 63 | dRMW-risk-092 | dMRW-risk-092 dRMW-risk-092 | Source cell contains identifier variants; canonical valid identifier selected | N/A | Labelling and Packaging | Shipping packaging remains intact during transit to customer | Shipping packaging does not remain intact during transit to customer | Product containers may fall out and not be provided to user | Fault | Delay to diagnosis (less than 10 days) | 2 | Packaging is not robust to withstand standard shipping to customer. | The product packaging is designed using standard Leica packaging specifications
LBN-dRMW-016 | 7 | 64 | dRMW-risk-093 | dMRW-risk-093 dRMW-risk-093 | Source cell contains identifier variants; canonical valid identifier selected | N/A | Assay performance | The product is run with the recommended HIER protocol | The product is not ran with the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Software allows user to select an enzyme pretreatment protocol | None Preservatives in formulation
LBN-dRMW-016 | 7 | 65 | dRMW-risk-094 | dRMW-risk-094 dRMW-risk-094 | Source cell contains identifier variants; canonical valid identifier selected | N/A | Storage and handling | The product shall maintain acceptable staining performance after transport and storage under the specified conditions | The product is frozen and subsequently thawed during transport, storage, or user handling | Use of product following a freeze–thaw temperature deviation resulting in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Inadequate temperature control during transport or storage, or accidental freezing by the user after receipt | Product labelling specifies the required storage conditions. Shipping packaging and distribution controls are intended to maintain the specified temperature range
LBN-dRMW-017 | 6 | 8 | dRMW-risk-001 | dRMW-risk-001 | Exact source identifier | N/A | CLC | The product is provided as ready-to-use slides requiring no processing prior to use | The product is not ready to use and requires processing prior to use | Additional processing required to achieve optimal staining and so when used as ready-to-use results in failure to meet acceptance criteria leading to a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | The product is designed to require processing prior to use | None
LBN-dRMW-017 | 6 | 9 | dRMW-risk-004 | dRMW-risk-004 | Exact source identifier | N/A | CLC | The product shall be able to be used directly from recommended storage | The product cannot be used directly from recommended storage | The product is loaded onto the BOND straight from storage and fails to meet acceptance criteria leading to a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | The product has not been designed to produce optimal staining when used directly from recommended storage | None
LBN-dRMW-017 | 6 | 10 | dRMW-risk-006 | dRMW-risk-006 | Exact source identifier | N/A | CLC | The FFPE cell cores must be presented in every section in the appropriate orientation | The FFPE cell cores are not presented in every section in the appropriate orientation | Product is loaded onto BOND-III and stained with HER2 BMA antibody, but fails to meet acceptance criteria as user scores the cell cores incorrectly based on expected cell core orientation shown in the Interpretation Guide | Normal | Delay to diagnosis (less than 10 days) | 2 | Cell cores not orientated correctly | None
LBN-dRMW-017 | 6 | 11 | dRMW-risk-007 | dRMW-risk-007 | Exact source identifier | N/A | CLC | The product contains three cell line cores which provide a HER2 staining range (high expression, low expression and negative) | The product does not contain three cell line cores which demonstrate a staining range (high expression, low expression and negative) | Product does not contain cores including a range of HER2 expression so is unable to act as a procedural control to detect a failed run | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative (HER2 null), False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Cell line control not designed with an appropriate range of HER2 expressing cores | None
LBN-dRMW-017 | 6 | 12 | dRMW-risk-009 | dRMW-risk-009 | Exact source identifier | N/A | CLC | The product is compatible with all components of the HER2 BMA Assay System | The product is not compatible with components of the HER2 BMA Assay System | Product is not compatible with HER2 BMA Assay System and does not stain correctly resulting in a false void run | Normal | Delay to diagnosis (more than 10 days) | 3 | Product has not been designed to be compatible with HER2 BMA Assay System | None
LBN-dRMW-017 | 6 | 13 | dRMW-risk-010 | dRMW-risk-010 | Exact source identifier | N/A | CLC | The Cell Line Control (CLC) functions as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The Cell Line Control (CLC) does not function as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The functionality demonstrated by the CLC is matrix‑dependent and may differ from that observed in tissue controls | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative (HER2 null), False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The CLC does not detect all analytical or biological failure modes | None
LBN-dRMW-017 | 6 | 14 | dRMW-risk-011 | dRMW-risk-011 | Exact source identifier | N/A | CLC | The product is compatible with clearing procedures using reagent grade alcohol, xylene, and a xylene/toluene-based mounting media | The Cell Line Control (CLC) functions as a procedural control within the HER2 BMA assay system to support interpretation of correct assay staining performance | The product is stained on the BOND-III, but is not compatible with general laboratory clearing and mounting procedures which results in difficulty in interpretation and a false void run | Normal | Delay to diagnosis (more than 10 days) | 3 | The product has not been designed to be compatible with laboratory clearing and mounting procedures or the workflow of the HER2 BMA Assay | None
LBN-dRMW-017 | 6 | 15 | dRMW-risk-013 | dRMW-risk-013 | Exact source identifier | N/A | Labelling & Packaging | The product IFU and interpretation guide explains the scoring system along with images of expected stained cell lines | The product IFU and interpretation guide does not explain the scoring system or contain images of expected stained cell lines | The product IFU and interpretation guide is inadequate and CLC failures are not identified | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | IFU/interpretation guide is not clear to aid in interpretation of the CLC | None
LBN-dRMW-017 | 6 | 16 | dRMW-risk-017 | dRMW-risk-017 | Exact source identifier | N/A | Stability | Each cell line core must maintain its HER2 expression level over its shelf life when stored at the recommended temperature (2-8°C) with minimum shelf life of 6 months for initial release | Each cell line does not maintain its HER2 expression level over the shelf life when stored at recommended temperature, and does not have a minimum shelf life of 6 months for initial release | Product is loaded onto BOND within shelf life but suboptimal staining occurs resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Unstable cell cores included in design of product | None
LBN-dRMW-017 | 6 | 17 | dRMW-risk-018 | dRMW-risk-018 | Exact source identifier | N/A | Transportation | The product maintains acceptable staining performance after transportation to the customer | The product does not maintain acceptable staining performance after transportation to the customer | Product is exposed to extreme conditions during transport and suboptimal staining occurs resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Extreme conditions could lead to instability and/or degradation | None
LBN-dRMW-017 | 6 | 18 | dRMW-risk-021 | dRMW-risk-021 | Exact source identifier | N/A | Product Performance | The product shall maintain HER2 expression levels when different lots of the CLC are used | The product does not maintain HER2 expression levels when different lots of the CLC are used | Product does not maintain HER2 expression levels across different lots of the CLC resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Failure to select robust cell line | None
LBN-dRMW-017 | 6 | 19 | dRMW-risk-021.1 | dRMW-risk-021.1 | Exact source identifier | N/A | Product Performance | The product shall maintain HER2 expression levels when different lots of antibody that have been qualified for use are used | The product does not maintain HER2 expression levels when different lots of the antibody are used | Product does not maintain HER2 expression levels across different lots of the antibody resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Failure to repeatably manufacture robust cell line | None
LBN-dRMW-017 | 6 | 20 | dRMW-risk-021.2 | dRMW-risk-021.2 | Exact source identifier | N/A | Product Performance | The product shall maintain HER2 expression levels when different lots of detection kit DS9800 that have been qualifieed for use are used | The product does not maintain HER2 expression levels when different lots of the detection kit DS9800 are used | Product does not maintain HER2 expression levels across different lots of the detection kit DS9800 resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Failure to select robust cell line | None
LBN-dRMW-017 | 6 | 21 | dRMW-risk-022 | dRMW-risk-022 | Exact source identifier | N/A | Product Performance | The product maintains HER2 expression levels when stained on multiple BOND-III instruments over multiple days | The product does not maintain HER2 expression levels when stained on multiple BOND-III instruments over multiple days | Product does not maintain HER2 expression levels when different BOND-III machines are used and on different days resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Failure to select robust cell line | None
LBN-dRMW-017 | 6 | 22 | dRMW-risk-028 | dRMW-risk-028 | Exact source identifier | N/A | Safety | The product shall present no hazards to the users that cannot be adequately controlled by good laboratory practices (GLP) | The product presents hazards to the users that cannot be adequately controlled by good laboratory practices | The product contains hazardous material that cannot be adequately controlled by GLP or product use may lead to user cutting themselves | Normal | User harm - mishandling hazardous product | 3 | User unaware of hazards | None
LBN-dRMW-017 | 6 | 23 | dRMW-risk-029 | dRMW-risk-029 | Exact source identifier | N/A | Labelling & Packaging | The product shall have labels on the immediate containers and outer packaging (if required) which will comply with local territory regulations | The product does not have labels on the immediate containers and or outer packaging which will comply with local territory regulations | The user is unable to determine the product and/or hazards as per their territory | Fault | User harm | 3 | Territories are not considered during label design and development | Regulatory hazard classifications considered
LBN-dRMW-017 | 6 | 24 | dRMW-risk-030 | dRMW-risk-030 | Exact source identifier | N/A | Labelling & Packaging | The product and immediate container must be labelled | The product and immediate container are not labelled with the appropriate information | The user cannot determine the product expiry and storage conditions and run the product out of specification, resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | The product label design does not contain appropriate information | None
LBN-dRMW-017 | 6 | 25 | dRMW-risk-031 | dRMW-risk-031 | Exact source identifier | N/A | Labelling & Packaging | The product and immediate container labels are legible indicating product, lot number and expiry date | The product and immediate container labels containing product, lot number and expiry date are not legible | The product and immediate container labels are not legible and the product is ran out of specification, resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | The product and immediate container labels are not legible | None
LBN-dRMW-017 | 6 | 26 | dRMW-risk-032 | dRMW-risk-032 | Exact source identifier | N/A | Labelling & Packaging | The outer and immediate packaging remains intact when stored at the recommended storage temperature | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | The outer and immediate packaging is damaged by storage at recommended temperature resulting in damage to the product meaning it cannot be used | Normal | Delay to diagnosis (less than 10 days) | 2 | Packaging design is not robust to withstand storage conditions | None
LBN-dRMW-017 | 6 | 27 | dRMW-risk-032.5 | dRMW-risk-032.5 | Exact source identifier | N/A | Labelling & Packaging | The outer and immediate packaging remains intact during shipping | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged or lost during shipping resulting in damage to the product meaning it cannot be used | Normal | Delay to diagnosis (less than 10 days) | 2 | Packaging design is not robust to withstand shipping conditions | None
LBN-dRMW-017 | 6 | 28 | dRMW-risk-033 | dRMW-risk-033 | Exact source identifier | N/A | Labelling & Packaging | The outer and immediate packaging must remain intact when stored at the recommended storage temperature | The outer and immediate packaging does not remain intact when stored at the recommended storage temperature | The outer and immediate packaging is damaged by storage at recommended temperature resulting in damage to the product which poses a cutting risk to the user | Normal | User harm | 3 | Packaging design is not robust to withstand storage conditions | None
LBN-dRMW-017 | 6 | 29 | dRMW-risk-033.5 | dRMW-risk-033.5 | Exact source identifier | N/A | Labelling & Packaging | The outer and immediate packaging remains intact during shipping | The outer and immediate packaging does not remain intact during shipping | The outer and immediate packaging is damaged during shipping resulting in damage which poses a cutting risk to the user | Normal | User harm | 3 | Packaging design is not robust to withstand shipping conditions | None
LBN-dRMW-017 | 6 | 30 | dRMW-risk-035 | dRMW-risk-035 | Exact source identifier | N/A | Labelling & Packaging | Product labelling shall instruct storage at 2-8°C | Product labelling does not instruct storage at 2-8°C | Suboptimal staining occurs as product has not been stored at recommended storage temperature which leads to a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Recommended storage temperature is not included within label design | None
LBN-dRMW-017 | 6 | 31 | dRMW-risk-039 | dRMW-risk-039 | Exact source identifier | N/A | Product Performance | The product is reproducible and has consistent appearance/ morphology between lots | The product is not reproducible and does not have consistent appearance/ morphology between lots | The product cannot be scored due to unacceptable cell line staining and morphology/appearance, resulting in a void run | Normal | Delay to diagnosis (less than 10 days) | 2 | Failure to select robust cell line | None
LBN-dRMW-017 | 6 | 32 | dRMW-risk-040 | dRMW-risk-040 | Exact source identifier | N/A | Assay Performance | The product must be able to be evaluated consistently by all pathologists | The product is not able to be evaluated consistently by different pathologists | CLC may not correctly be interpretated by all pathologists | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | The product has not been designed with clear acceptance criteria that gives consistent results when interpretated by all pathologists Robust cell line not selected Manufacturing procedure not robust | None
LBN-dRMW-017 | 6 | 33 | dRMW-risk-041 | dRMW-risk-041 | Exact source identifier | N/A | Labelling and Packaging | The product labels should have font and symbols that are human legible | The product labels do not have font and symbols that are human legible | The user cannot read the labelling font or the symbols on the product labels so cannot identify the product | Fault | Delay to diagnosis (less than 10 days) | 2 | The product labels designed with font and symbols too small to be read | None
LBN-dRMW-017 | 6 | 34 | dRMW-risk-042 | dRMW-risk-042 | Exact source identifier | N/A | Labelling and Packaging | The product labels should have font and symbols that are human legible | The product labels do not have font and symbols that are human legible | The user cannot read the labelling font or the symbols on the product labels so cannot identify the product | Fault | Delay to diagnosis (less than 10 days) | 2 | Font and symbols on the labels is not durable and can be rubbed off during use leading to illegible font/symbols | Label material/ink used for all current LBS cell line products
LBN-dRMW-017 | 6 | 35 | dRMW-risk-043 | dRMW-risk-043 | Exact source identifier | N/A | Labelling and Packaging | The product labels should have font and symbols that is legible to laboratory users | The product labels do not have font and symbols that is legible to laboratory users | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | Fault | Delay to diagnosis (less than 10 days) | 2 | Outer packaging labelling font and symbols are not designed to be durable for routine handling by user | Standard labelling and ink used
LBN-dRMW-017 | 6 | 36 | dRMW-risk-044 | dRMW-risk-044 | Exact source identifier | N/A | Labelling and Packaging | The product container labels shall remain adhered during the entire life of the product | The product container labels do not remain adhered during the entire life of the product | During the life of the product the labels become un-adhered to the product after exposure to reagents, rendering the user unable to identify the product/relevant safety labelling | Fault | User harm | 3 | The labels are not sufficiently adhesive to withstand the normal use of the product throughout its life | Label material used for all current LBS cell line control products
LBN-dRMW-017 | 6 | 37 | dRMW-risk-045 | dRMW-risk-045 | Exact source identifier | N/A | Labelling and Packaging | The product container label text and graphics shall be human legible after the product is exposed to reagents | The product container label text and graphics are not human legible after the product is exposed to reagents | The user cannot read the labelling font or the symbols on the container labels after exposure to reagents, rendering the user unable to identify relevant safety labelling | Fault | User harm | 3 | The label ink used is not sufficiently durable to withstand exposure to reagents | Label material/ink used for all current LBS cell line control products
LBN-dRMW-017 | 6 | 38 | dRMW-risk-046 | dRMW-risk-046 | Exact source identifier | N/A | Labelling and Packaging | The product packaging labels shall remain legible after exposure to normal storage and handling conditions | The product packaging labels are not legible after exposure to normal storage and handling conditions | After normal storage and handling the labels are not legible by the user, rendering the user unable to identify the product /relevant safety labelling | Normal | User harm | 3 | The label ink used is not sufficiently durable to withstand normal storage and handling | Label material/ink used for all current LBS cell line control products
LBN-dRMW-017 | 6 | 39 | dRMW-risk-047 | dRMW-risk-047 | Exact source identifier | SR-007 | Assay Performance | The product shall maintain HER2 expression levels across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The product does not maintain HER2 expression levels across the range of temperatures defined in the BOND-III user manual, 18–26 °C (64–79 °F) | The product does not maintain HER2 expression levels across the range of temperatures defined in the BOND-III user manual leading to a void run | Fault | Delay to diagnosis (less than 10 days) | 2 | The product is not robust enough to work across the range of BOND-III environmental conditions | BOND user manual specifies environmental conditions that the BOND-III can function satisfactorily under
LBN-dRMW-017 | 6 | 40 | dRMW-risk-048 | dRMW-risk-048 | Exact source identifier | SR-007 | Assay Performance | The product shall maintain HER2 expression levels across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The product does not maintain HER2 expression levels across the range of operating humidity defined in the BOND-III user manual, Operating humidity (noncondensing) 30 to 80% RH | The product does not maintain HER2 expression levels across the range of operating humidity's defined in the BOND user manual leading to a void run | Fault | Delay to diagnosis (less than 10 days) | 2 | The product is not robust enough to work across the range of BOND-III environmental conditions | BOND user manual specifies environmental conditions that the BOND-III can function satisfactorily under
LBN-dRMW-017 | 6 | 41 | dRMW-risk-049 | dRMW-risk-049 | Exact source identifier | SR-007 | Assay Performance | The product shall maintain HER2 expression levels across the range of altitudes defined in the BOND-III user manual, 0-1600m | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND III user manual, 0-1600m | The product does not maintain HER2 expression levels across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps leading to a void run | Normal | Delay to diagnosis (more than 10 days) | 3 | Customers at higher altitudes work in environments with lower atmospheric pressure, this reduces the boiling point which can affect the effectiveness of the HIER steps for unmasking the epitope as these steps are performed at 100°C with the HER2 BMA Assay | None
LBN-dRMW-017 | 6 | 42 | dRMW-risk-049.01 | dRMW-risk-049.01 | Exact source identifier | N/A | Assay Performance | The product shall maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude | The product does not maintain HER2 expression levels when ran with a user validated modified HIER protocol to mitigate testing at altitude | The product does not maintain HER2 expression levels when ran with a user valdiated modified HIER protocol and fails a run which should pass | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) False Negative, False Negative Result (HER2-low),False Negative Result (HER2-ultralow) | 4 | Customers at higher altitudes work in environments with lower atmospheric pressure, which may affect instrument performance; the HIER setting is unlocked, users may alter the programmed HIER conditions, leading to inconsistent antigen retrieval and variable staining results. | None BOND user manual specifies environmental conditions that the BOND-III can function satisfactorily under
LBN-dRMW-017 | 6 | 43 | dRMW-risk-049.1 | dRMW-risk-049.1 | Exact source identifier | N/A | Assay Performance | The product shall maintain HER2 expression levels and morphology when used above 1600m | The product does not maintain HER2 expression levels and morphology when used above 1600m | The product does maintain HER2 expression levels or morphology above 1600m leading to a valid run | Fault | Delay to diagnosis (more than 10 days) | 3 | Customers at higher altitudes work in environments with lower atmospheric pressure, this reduces the boiling point which can affect the effectiveness of the HIER steps for unmasking the epitope as these steps are performed at 100°C with the HER2 BMA Assay | None
LBN-dRMW-017 | 6 | 44 | dRMW-risk-050 | dRMW-risk-050 | Exact source identifier | N/A | Packaging and Labelling | The product container (slide mailer) remains intact during shipment to customer | The product container (slide mailer) does not remain intact during shipment to customer | Cell line cores may become contaminated or damaged and therefore cannot be used | Fault | Delay to diagnosis (less than 10 days) | 2 | Container is not sufficiently robust to withstand transport | Packaging selected is the same as all other LBS CLC products
LBN-dRMW-017 | 6 | 45 | dRMW-risk-051 | dRMW-risk-051 | Exact source identifier | N/A | Packaging and Labelling | The product container (slide mailer) is safe to handle | The product container (slide mailer) is not safe to handle | The user could be harmed by sharp edges of container | Normal | User harm- small cut | 2 | Container is composed of sharp and/or unsafe material | Packaging selected has no sharp edges and is the same as all other LBS CLC products.
LBN-dRMW-017 | 6 | 46 | dRMW-risk-052 | dRMW-risk-052 | Exact source identifier | N/A | Assay Performance | The product is run with the recommended HIER protocol | The product is not ran with the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | Normal | Delay to diagnosis (more than 10 days) | 3 | Software allows user to select an enzyme pretreatment protocol | None
LBN-euRMW-020 | 6 | 8 | euRMW-risk-003 | euRMW-risk-003 | Exact source identifier | N/A | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User embeds tissue specimens with a non-standard paraffin/embedding reagent. The tissue sections are stained sub-optimally | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended embedding technique | Standard ASCO-CAP/ CLSI I/LA28-A2 guidelines in place for tissue fixation/processing
LBN-euRMW-020 | 6 | 9 | euRMW-risk-005 | euRMW-risk-005 | Exact source identifier | N/A | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User prepares tissue sections outside that of the recommended thickness. The tissue sections are stained sub-optimally | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended tissue thickness | CLSI I/LA28-A2 and ASCO/CAP guidlines provide recommendations on best practice.
LBN-euRMW-020 | 6 | 10 | euRMW-risk-009 | euRMW-risk-009 | Exact source identifier | N/A | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User bakes tissue sections at different temperatures to that recommended resulting in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative result (HER2-ultralow) | 4 | User not aware of recommended tissue section baking conditions | CLSI I/LA28-A2 and ASCO/CAP guidlines provide recommendations on best practice.
LBN-euRMW-020 | 6 | 11 | euRMW-risk-011 | euRMW-risk-011 | Exact source identifier | N/A | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User bakes tissue sections at different time-frame to that recommended and results in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended tissue section baking conditions | CLSI I/LA28-A2 and ASCO/CAP guidlines provide recommendations on best practice.
LBN-euRMW-020 | 6 | 12 | euRMW-risk-011.1 | euRMW-risk-011.1 | Exact source identifier | N/A | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User bakes tissue sections off board and onboard the BOND-III instrument (double baked) | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended tissue baking conditions | CLSI I/LA28-A2 and ASCO/CAP guidlines provide recommendations on best practice.
LBN-euRMW-020 | 6 | 13 | euRMW-risk-011.2 | euRMW-risk-011.2 | Exact source identifier | N/A | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User does not bake tissue before use on the BOND-III instrument resulting in tissue lift | Fault | Delay to diagnosis (less than 10 days) | 2 | User not aware of recommended tissue baking conditions | CLSI I/LA28-A2 and ASCO/CAP guidlines provide recommendations on best practice.
LBN-euRMW-020 | 6 | 14 | euRMW-risk-011.5 | euRMW-risk-011.5 | Exact source identifier | N/A | CLC Prep | CLC preparation is carried out appropriately | CLC preparation is not carried out appropriately | User bakes the CLC resulting in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User is not aware that the CLC are provided pre-baked and ready to use | Aligns with industry best practice for commercially available slide products
LBN-euRMW-020 | 6 | 15 | euRMW-risk-013 | euRMW-risk-013 | Exact source identifier | N/A | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User prepares sections from non-breast carcinoma tissue resulting in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of intended use tissue | None
LBN-euRMW-020 | 6 | 16 | euRMW-risk-015 | euRMW-risk-015 | Exact source identifier | SR-019 | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User mounts tissue section outside of the recommended slide staining area, resulting in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended slide staining area | BOND User Manual recommends staining area and slide type (SRCM-027) Recommended slides contain annotated tissue placement area
LBN-euRMW-020 | 6 | 17 | euRMW-risk-017 | euRMW-risk-017 | Exact source identifier | SR-028 | Specimen handling | Tissue preparation is carried out appropriately | Tissue preparation is not carried out appropriately | User mounts tissue section onto non-recommended glass slides resulting in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended slide type | BOND User Manual recommends slide type (SRCM-027)
LBN-euRMW-020 | 6 | 18 | euRMW-risk-019 | euRMW-risk-019 | Exact source identifier | N/A | Scoring | Scoring of tissue slides results in correct result | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2 overexpression/HER2-low/HER2-ultralow case as HER2-null | Normal | False Negative Result (0 absent membrane staining) | 4 | Incorrect score assigned to patient case | Some scoring guidance is included in the ASCO/CAP guidelines but does not include all scoring categories
LBN-euRMW-020 | 6 | 19 | euRMW-risk-019.5 | euRMW-risk-019.5 | Exact source identifier | N/A | Scoring | Scoring of tissue slides results in correct result | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2 overexpression as a HER2-low or HER2-ultralow | Normal | False Negative (HER2-low), False Negative (HER2-ultralow) | 4 | Incorrect score assigned to patient case | Some scoring guidance is included in the ASCO/CAP guidelines but does not include all scoring categories
LBN-euRMW-020 | 6 | 20 | euRMW-risk-020 | euRMW-risk-020 | Exact source identifier | N/A | Scoring | Scoring of tissue slides results in correct result | Scoring of tissue slides results in incorrect result | Pathologist scores a HER2-null as HER2-low/HER2-ultralow/HER2 overexpression (including scoring artefactual/non-specific staining) | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) | 4 | Incorrect score assigned to patient case | Some scoring guidance is included in the ASCO/CAP guidelines but does not include all scoring categories
LBN-euRMW-020 | 6 | 21 | euRMW-risk-021 | euRMW-risk-021 | Exact source identifier | N/A | Scoring | Scoring of tissue slides results in correct result | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain not visible leading to error in scoring | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Weak haematoxylin counterstaining reduces contrast with the DAB chromogen, leading to inconsistent interpretation of the staining | None
LBN-euRMW-020 | 6 | 22 | euRMW-risk-022.1 | euRMW-risk-022.1 | Exact source identifier | N/A | Scoring | Scoring of tissue slides results in correct result | Scoring of tissue slides results in incorrect result | Haematoxylin counterstain too strong leading to error in scoring | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Strong haematoxylin counterstaining reduces contrast with the DAB chromogen, leading to inconsistent interpretation of the staining | None
LBN-euRMW-020 | 6 | 23 | euRMW-risk-023 | euRMW-risk-023 | Exact source identifier | N/A | Scoring | Scoring of tissue slides results in correct result | Scoring of tissue slides results in incorrect result | Pathologist scores using an alternative scoring method resulting in incorrect result | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Pathologist unaware of recommended scoring approach | Some scoring guidance is included in the ASCO/CAP guidelines, but does not include all scoring categories
LBN-euRMW-020 | 6 | 24 | euRMW-risk-024 | euRMW-risk-024 | Exact source identifier | N/A | Scoring | CLC slide is evaluated as intended | CLC slide is not evaluated as intended | Incorrect interpretation of the cell line controls may result in a run being unnecessarily voided | Fault | Delay to diagnosis (less than 10 days) | 2 | Pathologist not aware of how to interprerate the CLC slides | None
LBN-euRMW-020 | 6 | 25 | euRMW-risk-024.5 | euRMW-risk-024.5 | Exact source identifier | N/A | Scoring | CLC slide is evaluated as intended | CLC slide is not evaluated as intended | Incorrect interpretation of the CLC may result in a run being mistakenly passed | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Pathologist not aware of how to interprerate the CLC slides | None
LBN-euRMW-020 | 6 | 26 | euRMW-risk-032.5 | euRMW-risk-032.5 | Exact source identifier | N/A | Packaging and Labelling | Shipping packaging protects the product during transport | Shipping packaging does not protects the product during transport | Reagent container leaks during transport, exposing user to hazardous reagent | Normal | User harm | 3 | Reagent container damaged during transport | The product packaging follows standard Leica packaging specifications.
LBN-euRMW-020 | 6 | 27 | euRMW-risk-033.5 | euRMW-risk-033.5 | Exact source identifier | N/A | Packaging and Labelling | Outer packaging protects the product during opening | Outer packaging does not protect the product during opening | Reagent container leaks exposing user to hazardous reagent | Normal | User harm | 3 | User damages the reagent container while opening the outer packaging with a sharp instrument | The product packaging follows standard Leica packaging specifications, and includes an easy-tear feature
LBN-euRMW-020 | 6 | 28 | euRMW-risk-033.6 | euRMW-risk-033.6 | Exact source identifier | N/A | Packaging and Labelling | User can safely handle packaging | User can not safely handle packaging | User is cut whilst opening the packaging | Normal | User harm | 3 | User opens the outer packaging using a sharp instrument | The product packaging follows standard Leica packaging specifications, and includes an easy-tear feature
LBN-euRMW-020 | 6 | 29 | euRMW-risk-039 | euRMW-risk-039 | Exact source identifier | N/A | Packaging and Labelling | Reagent does not spill during opening of the reagent container | Reagent spills during opening of the reagent container | User opens reagent container and is exposed to hazardous reagent | Normal | User harm | 3 | Reagent container not easy to open | Existing BOND reagent containers are designed with an easy opening lid Appropriate PPE is required and routinely used in accordance with standard laboratory procedures
LBN-euRMW-020 | 6 | 30 | euRMW-risk-042 | euRMW-risk-042 | Exact source identifier | SR-002 | Labelling | The product (antibody) will not be used beyond the expiry date | The product (antibody) is used beyond the expiry date | The product is used beyond the expiry date leading to suboptimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative, False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User loads expired reagent onto the BOND-III instrument | BOND will not allow the use of expired reagents (SRCM-001)
LBN-euRMW-020 | 6 | 31 | euRMW-risk-043 | euRMW-risk-043 | Exact source identifier | N/A | Labelling | The product (CLC) will not be used beyond the expiry date | The product (CLC) is used beyond the expiry date | The user uses the product (CLC) beyond the expiry date and suboptimal staining leads to a incorrect void run | Fault | Delay to diagnosis (less than 10 days) | 2 | User unaware of expiry date | None
LBN-euRMW-020 | 6 | 32 | euRMW-risk-044 | euRMW-risk-044 | Exact source identifier | N/A | Health and Safety | Correct PPE is worn when using the product | Incorrect /no PPE is worn when using the product | The user comes into contact with the reagent which contains 0.035% 2-methylisothiazol-3(2H)-one | Normal | User harm | 3 | User is unaware of PPE requirements | Appropriate PPE is required and routinely used in accordance with standard laboratory procedures
LBN-euRMW-020 | 6 | 33 | euRMW-risk-046 | euRMW-risk-046 | Exact source identifier | N/A | Health and Safety | User disposes of reagents appropriately | User does not dispose of reagents appropriately | The user does not dispose of the product appropriately. The component parts (0.035% 2-methylisothiazol-3(2H)-one) are dangerous for the environment | Normal | Environmental Harm | 3 | User not aware of recommended disposal | Federal/State or local regulations are in place for disposal of potentially toxic components
LBN-euRMW-020 | 6 | 34 | euRMW-risk-047 | euRMW-risk-047 | Exact source identifier | N/A | Health and Safety | User disposes of glass sides appropriately | User does not dispose of glass sides appropriately | Glass slides not properly disposed and may cause injury | Fault | User harm | 2 | User unaware of recommended disposal | Federal/State or local regulations are in place for disposal of sharps and human tissue
LBN-euRMW-020 | 6 | 35 | euRMW-risk-048 | euRMW-risk-048 | Exact source identifier | SR-024 | Customer usage | The BOND run is set up correctly | The BOND run is not set up correctly | Covertiles misplaced on the slide leading to sub-optimal staining | Fault | False Negative, False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User error during run set up | BOND User Manual gives guidance on correct covertile use (SRCM-029) SSA is moulded to guide correct placement
LBN-euRMW-020 | 6 | 36 | euRMW-risk-049 | euRMW-risk-049 | Exact source identifier | SR-024 | Customer usage | The BOND run is set up correctly | The BOND run is not set up correctly | Covertiles are not placed on top of the slides leading to sub-optimal staining | Fault | False Negative, False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User error during run set up | BOND User Manual gives guidance on correct covertile use (SRCM-029) Staining would not be detectable on the tissue section
LBN-euRMW-020 | 6 | 37 | euRMW-risk-050 | euRMW-risk-050 | Exact source identifier | SR-008 | Customer usage | The BOND run is set up correctly | The BOND run is not set up correctly | Unclean covertiles are used for staining leading to background staining and/or sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative, False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User error in covertile cleaning or BOND User Maintenance Guidelines not followed | BOND User Manual gives guidance on covertile use and cleaning (SRCM-014)
LBN-euRMW-020 | 6 | 38 | euRMW-risk-052 | euRMW-risk-052 | Exact source identifier | SR-009 | Customer usage | The BOND run is set up correctly | The BOND run is not set up correctly | Insufficient BOND ancillary reagents loaded onto BOND leading to BOND run error or delay | Fault | Delay to diagnosis (less than 10 days) | 2 | User error when setting up BOND run, ancillary reagents not checked as per BOND User Manual | BOND error and/or GUI notification will notify user to replenish any ancillary reagents required if the instrument is loaded with insufficient ancillary reagents (SRCM-016) BOND run will not start without sufficient reagent (SRCM-015)
LBN-euRMW-020 | 6 | 39 | euRMW-risk-053 | euRMW-risk-053 | Exact source identifier | N/A | Customer usage | The user coverslips using an appropriate mountant | The user coverslips using an non-appropriate mountant | User coverslips using non-recommended mountant, leading to sub-optimal visualization of staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative, False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended mountant | None
LBN-euRMW-020 | 6 | 40 | euRMW-risk-055 | euRMW-risk-055 | Exact source identifier | N/A | Customer usage | The product is used for staining on freshly cut FFPE tissue sections | The product is used for staining on aged FFPE tissue sections | FFPE tissue sections used beyond antigen stability leading to sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative, False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended cut section stability | ASCO-CAP guidelines recommend staining within 6 weeks
LBN-euRMW-020 | 6 | 41 | euRMW-risk-058 | euRMW-risk-058 | Exact source identifier | SR-010 | BOND Maintenance | The BOND-III instrument is maintained as per IFU and BOND User Manual | The BOND-III instrument is not maintained as per IFU and BOND User Manual | The BOND-III instrument is not cleaned and mould/contaminants in tubing lead to sub-optimal staining. | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of BOND cleaning protocol or BOND User Maintenance Guidelines | BOND User Manual contains complete details on maintenance and cleaning requirements (SRC-018)
LBN-euRMW-020 | 6 | 42 | euRMW-risk-064 | euRMW-risk-064 | Exact source identifier | SR-011 | Customer usage | A dispense volume of 150 µL is used during application of the product to tissue sections | A dispense volume of 150 µL is not used during application of the product to tissue sections | Slide is set up as 100 µL dispense and the user attempts to use it on a BOND-III instrument | Fault | Delay to diagnosis (less than 10 days) | 2 | The user is unaware that 100 µL dispense is not compatible with the assay | BOND-III can only dispense 150 µL. A BOND error message will be displayed and the run will not start if the 100 µL is selected (SRCM-019).
LBN-euRMW-020 | 6 | 43 | euRMW-risk-066 | euRMW-risk-066 | Exact source identifier | N/A | Customer usage | The recommended detection kit is used with the product to perform staining | An alternative detection kit to that recommended is used with the product to perform staining | Slides stained using alternative detection kit to the recommended DS9800 leading to sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of the recommended detection kit | None
LBN-euRMW-020 | 6 | 44 | euRMW-risk-068 | euRMW-risk-068 | Exact source identifier | SR-022 | Customer usage | The product is used for automated staining on the BOND-III staining platform | The product is not used for automated staining on the BOND-III staining platform | The product is used on an alternative automated staining platform to the BOND-III leading to sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User is unaware that the product is for use on the BOND-III instrument and loads slides on a non-recommended BOND staining platform | None
LBN-euRMW-020 | 6 | 45 | euRMW-risk-070 | euRMW-risk-070 | Exact source identifier | SR-006 | Customer usage | The product is run on the BOND-III using the recommended BOND software version | The product is ran on the BOND-III using a non recommended BOND software version | The BOND-III is on a non-recommended software version leading to sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User loads the product onto a BOND-III using an older software version than what is recommended | None
LBN-euRMW-020 | 6 | 46 | euRMW-risk-072 | euRMW-risk-072 | Exact source identifier | N/A | Customer usage | Assay is run alongside CLCs | Assay is not run alongside CLCs | CLCs are not included in run and sub-optimal staining cannot be detected CLCs are not included in run and certain changes in procedural staining performance may not be identified | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of or does not follow recommendation of use of CLCs as procedural control | None
LBN-euRMW-020 | 6 | 47 | euRMW-risk-075 | euRMW-risk-075 | Exact source identifier | N/A | Customer usage | Product is run alongside recommended tissue controls | Product is not run alongside recommended tissue controls | Tissue controls are not included in run and sub-optimal staining cannot be detected | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of or does not follow recommendation of use of tissue controls | None Established practice and professional guidelines (CLSI I/LA28‑A2 and ASCO CAP guidelines) requiring tissue controls to confirm run validity
LBN-euRMW-020 | 6 | 48 | euRMW-risk-078 | euRMW-risk-078 | Exact source identifier | SR-023 | Customer usage | Product is run with recommended staining protocol | Product is not run with recommended staining protocol | User selects a non-recommended staining protocol leading to sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), FALSE Negative Result (HER2-ultralow) | 4 | User unaware of recommended staining protocol | None
LBN-euRMW-020 | 6 | 49 | euRMW-risk-080 | euRMW-risk-080 | Exact source identifier | N/A | Customer usage | Product is run with recommended negative reagent control | Product is not run with recommended negative reagent control | Negative reagent controls are not included in run and sub-optimal staining cannot be detected | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of or does not follow recommendation of use of tissue controls | None
LBN-euRMW-020 | 6 | 50 | euRMW-risk-083 | euRMW-risk-083 | Exact source identifier | N/A | Storage and Handling | The product is stored within the recommended temperature range | The product is not stored within the recommended temperature range | User does not store product at recommended temperature range (2-8°C). Following storage the product is used on the BOND-III and results in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of or does not follow recommended storage conditions (2-8°C) | None
LBN-euRMW-020 | 6 | 51 | euRMW-risk-086 | euRMW-risk-086 | Exact source identifier | N/A | Specimen handling | Slides are dehydrated and cleared as per standard IHC techniques | Slides are not dehydrated and cleared as per standard IHC techniques | User does not appropriately dehydrate and clear the sections resulting in bubbles/debris on slides that interfere with slide interpretation | Fault | Delay to diagnosis (less than 10 days) | 2 | User not aware of recommened dehydration and clearing procedures | Assay designed to use standard IHC alcohol-xylene series dehydrating and clearing techniques as detailed in CLSI I/LA28-A2
LBN-euRMW-020 | 6 | 52 | euRMW-risk-089 | euRMW-risk-089 | Exact source identifier | SR-032 | Customer usage | Product is run with recommended HIER protocol | Product is not run with recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | User not aware of recommended HIER protocol | None
LBN-euRMW-020 | 6 | 53 | euRMW-risk-091 | euRMW-risk-091 | Exact source identifier | N/A | HER2 Antibody | The product maintains acceptable staining performance over the life of the antibody on the BOND III instrument | The product does not maintain acceptable staining performance over the life of the antibody on the BOND III instrument | User leaves product on the BOND-III for extended periods of time between use, leading to antibody degradation and evaporation | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow) | 4 | User unaware of appropriate storage conditions when product is not in use | None
LBN-euRMW-020 | 6 | 54 | euRMW-risk-092 | euRMW-risk-092 | Exact source identifier | SR-016 | HER2 Antibody | The product will be ready to use and does not require any mixing or shaking before use | The product is not ready to use and requires mixing or shaking before use | User shakes the product prior to use introducing bubbles, resulting in sub-optimal staining | Fault | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Bubbles interfere with dispense of reagent | SCRM-026
LBN-euRMW-020 | 6 | 55 | euRMW-risk-093 | euRMW-risk-093 | Exact source identifier | SR-025 | BOND Set Up | The HER2 BMA CDx Assay is able to be used with the BOND-III delayed start function | Use of the delayed start function for extended periods of time when running the HER2 BMA CDx Assay | User uses BOND-III delayed start function which results in sub-optimal staining | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Use of the BOND‑III delayed‑start function results in reduced stability of the HER2 BMA CDx Assay while on the instrument | BOND restricts delayed start function to no more than 7 days. (SRCM-031) BOND User Manual states maximum time limit
LBN-euRMW-020 | 6 | 56 | euRMW-risk-096 | euRMW-risk-096 | Exact source identifier | SR-027 | BOND Set Up | The BOND run completes successfully and the user unloads the SSA in a timely manner | The BOND run completes successfully and the user does not unload the SSA in a timely manner | The user does not unload the SSA, the slides is continually hydrated and staining becomes sub-optimal over time | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | Continued hydration of slide weakens staining signal | BOND GUI notifies users when processing of slides is complete, so that users can remove stained slides. (SRCM-030) BOND User Manual states the prompt removal of slide after run.
LBN-euRMW-020 | 6 | 57 | euRMW-risk-097 | euRMW-risk-097 | Exact source identifier | SR-032 | Running of HER2 BMA CDx Assay (open HIER protocol) | User uses the recommended HIER protocol | User does not use the recommended HIER protocol | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining | Normal | False Positive (Overexpression), False Positive (HER2-low), False Positive (HER2-ultralow), False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | HIER open so user can make changes to the protocol User select a different HIER protocol that is not internally optimized or validated for use with the product | Validated HIER protocol selected as default when scanning antibody onto BOND-III instrument (SCRM-025) User has to manually change HIER protocol Standard lab practice requires validation of protocol changes
LBN-euRMW-020 | 6 | 58 | euRMW-risk-098 | euRMW-risk-098 | Exact source identifier | SR-021 | Running of HER2 BMA CDx Assay (open HIER protocol) | User runs assay at altitudes less than 1600m | User runs assay at altitudes greater than 1600m | Staining in tissue samples differs to the staining that would have resulted if ran at altitudes below 1600m Epitope retrieval failure leads to sub optimal staining or patchy staining | Normal | False Negative Result (HER2-null), False Negative Result (HER2-low), False Negative Result (HER2-ultralow) | 4 | High altitude environment has lower atmospheric pressure and thus the boiling point of water is lower than 100°C The lower boiling point impacts the staining/tissue as the recommended HIER protocol is set at 100°C, boiling off of the ER solution causes tissue to be dried or epitope retrieval failure | BOND User Manual states that BOND-III instrument is only for use up to 1600m (SRCM-004)

## PR Register (hidden)
ProductFamily | ProductRequirementID | ProductRequirementText | AcceptanceCriteria | AssessmentDetails | ReferenceMaterials | SourceDocument | SourceRevision | SourceLocation | ExtractionNotes
DP52 | PR.01 | The Product must be specific for the HER2 protein. | WESTERN BLOT Band for The Product must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% of staining on Normal Tissues must meet the staining described in DP52-SAC-001. ≥95% of staining on Abnormal Tissues must meet the staining described in DP52-SAC-001. | DP52-VEP-011 - Analytical Specificity and Sensitivity Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | CLSI EP17-A2:2012 Evaluation of Detection Capability for Clinical Laboratory Measurement Procedures; CLSI I/LA28-A2 Quality Assurance for Design Control and Implementation of Immunohistochemistry Assays; Guidance for Submission of Immunohistochemistry Applications to the FDA | DP52-VEP-001 | 1 | Page 12-13
DP52 | PR.02 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues. | DP52-VEP-013 - System Compatibility Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 14
DP52 | PR.03 | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | DP52-VEP‑005 - Prevalence Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 14-15
DP52 | PR.04 | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining). | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) and describes how staining of the HER2 BMA CDx Assay can be interpreted. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | FDA guidance for IVD Device labelling; EU Labelling requirements for IVD; ISO 18113‑1 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements; ISO 18113‑2 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; EN ISO 15223‑1 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; EN ISO 14971 – Medical devices — Application of risk management to medical devices | DP52-VEP-001 | 1 | Page 15-17
DP52 | PR.05 | The Product must be compatible with the HER2 BMA Cell Line Control, which will serve as the procedural control for the HER2 BMA Assay. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | DP52-VEP-008 - Robustness Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 18
DP52 | PR.06 | The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | DP52-VEP-010 - Sample Characterization Verification Plan for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | ASCO-CAP Guidelines; ESMO Guidelines | DP52-VEP-001 | 1 | Page 18-19
DP52 | PR.07 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness) as determined by verification studies. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | BS EN ISO 15223‑1:2021+A1:2025 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; BS EN ISO 18113‑2:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; BS EN ISO 18113‑1:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements | DP52-VEP-001 | 1 | Page 19-21
DP52 | PR.08 | The Product Instructions for Use (IFU) document must state any limitations relating to tissue fixation. | Product Instructions for Use (IFU) document contains details of any limitations related to tissue fixation (e.g., fixative type, fixation duration, over or under fixation effects) as determined by verification studies. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | BS EN ISO 15223‑1:2021+A1:2025 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; BS EN ISO 18113‑2:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; BS EN ISO 18113‑1:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements | DP52-VEP-001 | 1 | Page 21-23
DP52 | PR.10 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim. | DP52-VEP-012 - Reagent Stability Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | BS EN ISO 23640:2015 – In vitro diagnostic medical devices — Evaluation of stability of in vitro diagnostic reagents; CLSI EP25-Ed2:2023 – Evaluation of Stability of In Vitro Medical Laboratory Test Reagents | DP52-VEP-001 | 1 | Page 23-24
DP52 | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | DP52-VEP-012 - Reagent Stability Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | BS EN ISO 23640:2015 – In vitro diagnostic medical devices — Evaluation of stability of in vitro diagnostic reagents; CLSI EP25-Ed2:2023 – Evaluation of Stability of In Vitro Medical Laboratory Test Reagents | DP52-VEP-001 | 1 | Page 24-25
DP52 | PR.12 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | DP52-VEP-013 System Compatibility Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 25-26
DP52 | PR.13 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | DP52-VEP-013 - System Compatibility Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 26
DP52 | PR.14 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | RP101-IUVP-001 - NPD Overarching In-Use Verification Plan; RP101-IUVR-001 - NPD Overarching In-Use Verification Report; See Section 4.3.2 for assessment of existing data. | N/A | DP52-VEP-001 | 1 | Page 27
DP52 | PR.15 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | RP101-IUVP-001 - NPD Overarching In-Use Verification plan; RP101-IUVR-001 - NPD Overarching In-Use Verification Report; See Section 4.3.2 for existing data applicable to DP52. | N/A | DP52-VEP-001 | 1 | Page 27-28
DP52 | PR.16 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum six months shelf life with a minimum two months safety margin). | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | DP52-VEP-012 - Reagent Stability Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | BS EN ISO 23640:2015 – In vitro diagnostic medical devices — Evaluation of stability of in vitro diagnostic reagents; CLSI EP25-Ed2:2023 – Evaluation of Stability of In Vitro Medical Laboratory Test Reagents | DP52-VEP-001 | 1 | Page 28-29
DP52 | PR.17 | The Product will have a minimum cut section stability of 30 days. | 100% of all HER2 test slides stored and aged at 2-8°C stained with The Product must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated cut section stability claim. | DP52-VEP-002 - Cut Section Stability Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | BS EN ISO 23640:2015 – In vitro diagnostic medical devices — Evaluation of stability of in vitro diagnostic reagents; CLSI EP25-Ed2:2023 – Evaluation of Stability of In Vitro Medical Laboratory Test Reagents | DP52-VEP-001 | 1 | Page 29-30
DP52 | PR.18 | The Product labelling will detail known limitations of formalin fixed, paraffin-embedded (FFPE) cut section stability. | Product Instructions for Use (IFU) document specifies any known limitations for cut tissue sections during the stability period as determined by verification testing. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | BS EN ISO 15223‑1:2021+A1:2025 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; BS EN ISO 18113‑2:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; BS EN ISO 18113‑1:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements | DP52-VEP-001 | 1 | Page 30-32
DP52 | PR.19 | The Product must be correctly identified by the BOND-III fully automated system. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | DP52-VEP-013 - System Compatibility Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 32-33
DP52 | PR.20 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings. | DP52-VEP-013 - System Compatibility Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 33
DP52 | PR.21 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | DP52-VEP-013 - System Compatibility Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 33-34
DP52 | PR.22 | The Product Instructions for Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | Product Instructions for Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | BS EN ISO 15223‑1:2021+A1:2025 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; BS EN ISO 18113‑2:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; BS EN ISO 18113‑1:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements | DP52-VEP-001 | 1 | Page 34-36
DP52 | PR.23 | The Product, when used within the HER2 BMA CDx Assay System, must generate staining that can be interpreted in the same way by different pathologists (inter-reader precision). | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). | LBS-CA-PTL-064 Swift Reproducibility Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 36-37
DP52 | PR.24 | The Product, when used within the HER2 BMA CDx Assay System, must generate staining that can be interpreted in the same way by the same pathologist over time (intra-reader precision). | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). | LBS-CA-PTL-064 Swift Reproducibility Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 37-38
DP52 | PR.25 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of HER2 BMA Antibody CDx on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. This acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | DP52-VEP-007 - Precision Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | I/LA28-A2 – Quality Assurance for Design Control and Implementation of Immunohistochemistry Assays; Approved Guideline—Second Edition; CLSI EP05:2025 – Evaluation of Precision Performance of Quantitative Measurement Methods; CLSI EP15-A3:2014 – User Verification of Precision and Estimation of Bias; Approved Guideline—Third Edition | DP52-VEP-001 | 1 | Page 38-40
DP52 | PR.26 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | DP52-VEP-007 - Precision Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | I/LA28-A2 – Quality Assurance for Design Control and Implementation of Immunohistochemistry Assays; Approved Guideline—Second Edition; CLSI EP05:2025 – Evaluation of Precision Performance of Quantitative Measurement Methods; CLSI EP15-A3:2014 – User Verification of Precision and Estimation of Bias; Approved Guideline—Third Edition | DP52-VEP-001 | 1 | Page 40-42
DP52 | PR.27 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different BOND-III systems. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | DP52-VEP-007 - Precision Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | I/LA28-A2 – Quality Assurance for Design Control and Implementation of Immunohistochemistry Assays; Approved Guideline—Second Edition; CLSI EP05:2025 – Evaluation of Precision Performance of Quantitative Measurement Methods; CLSI EP15-A3:2014 – User Verification of Precision and Estimation of Bias; Approved Guideline—Third Edition | DP52-VEP-001 | 1 | Page 42-43
DP52 | PR.28 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | Tissue cases must demonstrate Acceptable Staining and Consistent Staining across three different SSAs on a BOND-III Instrument run batch. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | DP52-VEP-007 - Precision Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | I/LA28-A2 – Quality Assurance for Design Control and Implementation of Immunohistochemistry Assays; Approved Guideline—Second Edition; CLSI EP05:2025 – Evaluation of Precision Performance of Quantitative Measurement Methods; CLSI EP15-A3:2014 – User Verification of Precision and Estimation of Bias; Approved Guideline—Third Edition | DP52-VEP-001 | 1 | Page 44-45
DP52 | PR.29 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining within BOND Runs (intra-run repeatability). | For each of the conditions listed below, a minimum of 90% of tissue cases must demonstrate Acceptable Staining and Consistent Staining across the same SSA on a BOND-III Instrument. Negative Percent Agreement (NPA), Positive Percent Agreement (PPA), and Overall Percent Agreement (OPA) will be evaluated. The acceptance criteria require a point estimate of ≥90% agreement, the percentage agreement rate to be ≥85% at the lower bound 95% CI. | DP52-VEP-007 - Precision Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | I/LA28-A2 – Quality Assurance for Design Control and Implementation of Immunohistochemistry Assays; Approved Guideline—Second Edition; CLSI EP05:2025 – Evaluation of Precision Performance of Quantitative Measurement Methods; CLSI EP15-A3:2014 – User Verification of Precision and Estimation of Bias; Approved Guideline—Third Edition | DP52-VEP-001 | 1 | Page 46-47 | [NEEDS INPUT] Acceptance Criteria says "conditions listed below", but no conditions were present in the extracted table text.
DP52 | PR.30 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). | LBS-CA-PTL-064.R01 Swift Reproducibility Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 47-48
DP52 | PR.31 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining in different laboratories (inter-site reproducibility). | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for APA and ANA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Average positive agreement (APA), average negative agreement (ANA), and average overall percent agreements (OPA), along with corresponding 2-sided 95% confidence intervals (CIs), for each of the 2 by 2 analyses (HER2 overexpression and HER2 low/ultralow expression) will be calculated for each of the 4 sub-studies (inter-site, inter-day, inter-pathologist and intra-pathologist). Secondary endpoint: Agreement measures for each sub-study described above (primary endpoint) using 4 categories (0, 1+, 2+, and 3+) of HER2 IHC scores. All agreement measures will be calculated along with corresponding 2-sided 95% confidence intervals (CIs). | LBS-CA-PTL-064 Swift Reproducibility Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 49-50
DP52 | PR.32 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (Dako HercepTest ™ ) for HER2 overexpression. | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). | LBS-CA-PTL-063 Swift IHC Method Comparison Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 50-51
DP52 | PR.33 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (VENTANA® PATHWAY® anti-HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody) for HER2-low and HER2-ultralow. | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). | LBS-CA-PTL-063 Swift IHC Method Comparison Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 51-52
DP52 | PR.34 | The Product, when used within the HER2 BMA CDx Assay System, must generate comparable staining with the comparator device (VENTANA® PATHWAY® anti-HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody) for HER2 overexpression. | Performance will be deemed acceptable if the applicable primary and secondary endpoints meet the following criteria: agreement rate shall be ≥90% for PPA and NPA with each respective lower bounds of the 2-sided 95% CIs ≥ 85%. Primary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator under binary (positive/negative) classification at four cut-offs: HER2 ultralow (vs. 4B5 only): negative = 0 null; positive = 0 ultralow/1+/2+/3+ HER2 low (vs. 4B5 + HercepTest): negative = 0 null/ultralow; positive = 1+/2+/3+ HER2 overexpression (vs. 4B5 + HercepTest): negative = 0 null/ultralow/1+; positive = 2+/3+ HER2 overexpression, 3+ cut-off (vs. 4B5 + HercepTest): negative ≤2+; positive = 3+ Secondary endpoint: Concordance (OPA/PPA/NPA) between the HER2 BMA CDx Assay and each comparator (4B5 + HercepTest) under 4-category classification (0/1+/2+/3+). | LBS-CA-PTL-063.R02 Swift IHC Method Comparison Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 52-53
DP52 | PR.35 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96 °C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | DP52/DP61-VEP-001 – Environmental Verification Study Protocol for the HER2 BMA CDx Assay | N/A | DP52-VEP-001 | 1 | Page 54
DP52 | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating temperature range of 18-26°C (64-79°F on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. | DP52/DP61-VEP-014 – Temperature and Humidity Verification Study Protocol for the HER2 BMA CDx Assay | N/A | DP52-VEP-001 | 1 | Page 54-55 | [NEEDS INPUT] Extracted text contains an unmatched parenthesis after "64-79°F" and the unclear term "*IHC P"; preserved exactly.
DP52 | PR.37 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III humidity specifications. | Analytical evidence shall show that the instrument performs with acceptable staining quality across the operating humidity (noncondensing) 30 to 80% RH range on the BOND instruments. At minimum, this includes achieving a 95% pass rate for acceptable staining for markers using staining protocols functionally representative of *IHC P used by The Product. | DP52/DP61-VEP-014 – Temperature and Humidity Verification Study Protocol for the HER2 BMA CDx Assay | N/A | DP52-VEP-001 | 1 | Page 55 | [NEEDS INPUT] Extracted text contains the unclear term "*IHC P"; preserved exactly.
DP52 | PR.38 | The Product Instructions for Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | Product Instructions for Use (IFU) states any limitations on product use related to altitude, temperature, and humidity from verification testing. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | BS EN ISO 15223‑1:2021+A1:2025 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; BS EN ISO 18113‑2:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; BS EN ISO 18113‑1:2024 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements | DP52-VEP-001 | 1 | Page 55-57
DP52 | PR.39 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | DP52-VEP-013 – System Compatibility Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 58
DP52 | PR.40 | The Product can be run within the same BOND Run Batch as other BOND Detection systems and staining protocol combinations. | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | DP52-VEP-013 – System Compatibility Verification Protocol for the bond Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 58-59
DP52 | PR.41 | The Product must generate Acceptable Staining when there is a delay of up to seven days between placing The Product on the BOND-III and commencement of staining. | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | DP52-VEP-008 - Robustness Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 59
DP52 | PR.42 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining when there is a delay of up to 24 hours between completion of staining and removal. | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | DP52-VEP-008 - Robustness Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 59-60
DP52 | PR.43 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification of Chemical Reagents). | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | FDA guidance for IVD Device labelling; EU Labelling requirements for IVD; ISO 18113‑1 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements; ISO 18113‑2 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; EN ISO 15223‑1 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; EN ISO 14971 – Medical devices — Application of risk management to medical devices | DP52-VEP-001 | 1 | Page 60-63
DP52 | PR.44 | The symbols and text included on The Product labelling must be human legible. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | FDA guidance for IVD Device labelling; EU Labelling requirements for IVD; ISO 18113-1; ISO 18113-2; EN ISO 15223-1; EN ISO 14971 | DP52-VEP-001 | 1 | Page 63
DP52 | PR.45 | The Product when expired, will not allow a staining run to commence. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | DP52-VEP-013 – System Compatibility Verification Protocol for the bond Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 64
DP52 | PR.46 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | DP52-VEP-003 - Labelling Verification Protocol for the bond Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 64
DP52 | PR.47 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging remains serviceable following 46 thermal cycles designed to mimic on BOND and storage conditions (i.e. 7hr at maximum expected temperature of 34ºC, storage at 2-8ºC). AND One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. | RP101-OPVP-001 Overarching Packaging Verification Plan; RP101-OPVR-001 Overarching Packaging Verification Report; Further details of the assessments can be found in following protocols: RP101-PLVSP-002 Product Packaging and Labeling Transport and In Use Stability Verification Protocol; RP101-PLVSR-002 Product Packaging and Labeling Transport and In Use Stability Verification Report; RP101- PLVSP-003 Product Packaging Robustness Verification Protocol; RP101- PLVSR-003 Product Packaging Robustness Verification Report; See Section 4.3 for existing data applicable to DP52. | N/A | DP52-VEP-001 | 1 | Page 65-66
DP52 | PR.48 | The Outer Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | One laboratory user assessment results in 100% of 32 samples: • The product outer (pouch), shrink sleeve, and immediate (container) packaging and labelling remain serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user confirms that after the product is dropped once from 1m height on top, bottom, left or right side: Immediate and outer product packaging remains serviceable with no evidence of reagent container leakage and can be loaded into a BOND reagent rack. | RP101-OPVP-001 Overarching Packaging Verification Plan; RP101-OPVR-001 Overarching Packaging Verification Report; Further details of the assessments can be found in following protocols: RP101-PLVSP-002 Product Packaging and Labeling Transport and In Use Stability Verification Protocol; RP101-PLVSR-002 Product Packaging and Labeling Transport and In Use Stability Verification Report; RP101- PLVSP-003 Product Packaging Robustness Verification Protocol; RP101- PLVSR-003 Product Packaging Robustness Verification Report; See Section 4.3 for existing data applicable to DP52. | N/A | DP52-VEP-001 | 1 | Page 67-68
DP52 | PR.49 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | Assessment results in 100% samples demonstrating that packaging of products contained within the different shipping boxes being tested (ZSV003, ZCB033, ZCB292, ZCB172, ZCB035, and ZCB088) is undamaged and serviceable via visual inspection when the shipping box is: • dropped from a minimum height of 0.8m 2 , onto a hard surface, on: o the face of the packaging with the largest surface area o one corner of the shipping package pointing downwards • Boxes are stacked following EU regulations along with other box types and shipments (max height of 150cm). | RP101-SPVP-001 Overarching Verification Plan for Shipping Packaging; RP101-SPVR-001 Overarching Verification Report for Shipping Packaging; See Section 4.3 for existing data applicable to DP52. | N/A | DP52-VEP-001 | 1 | Page 69 | [NEEDS INPUT] "0.8m 2 ," appears to include a separated footnote marker or extraction artifact; preserved exactly.
DP52 | PR.50 | The barcodes included on The Product reagent container label must remain fully readable by the BOND-III fully automated system following routine shipping and storage. | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. | Doc ID: RP101-OPVP-001 Overarching Packaging Verification Plan; Further details of the assessments can be found in following protocols: RP101-PLVSP-002 Product Packaging and Labeling Transport and In Use Stability Verification Protocol; RP101- PLVSP-004 Product Label Chemical Resistance Verification Study Protocol | N/A | DP52-VEP-001 | 1 | Page 69-71
DP52 | PR.51 | The barcodes, symbols and text included on The Product reagent container label must remain human legible following routine shipping and storage. | One laboratory user assessment results in 100% of 32 samples: • The product labelling remains serviceable after undergoing a simulated transportation test designed to mimic the conditions that the product might conceivably experience during transport (2 days at 30°C, at least 1 day at 2-8°C and 2 days at 40°C). AND One laboratory user assessment results in 100% of 32 samples: • For 7 mL products, with expected customer use of up to 46 tests, the product immediate (reagent container) packaging label remains serviceable following 46 thermal cycles. AND • Reagent is incubated at room temperature (within the range of 19-23°C) on the immediate packaging label for at least 10mins, then removed. Following reagent removal, one assessor confirms the readability of the product immediate packaging label following chemical exposure (as per the assessment in the test protocol) and successful scan by a BOND console. | Doc ID: RP101-OPVP-001 Overarching Packaging Verification Plan; Further details of the assessments can be found in following protocols: RP101-PLVSP-002 Product Packaging and Labeling Transport and In Use Stability Verification Protocol; RP101- PLVSP-004 Product Label Chemical Resistance Verification Study Protocol | N/A | DP52-VEP-001 | 1 | Page 71-72
DP52 | PR.52 | The Product shall include compliance with UDI specifications. | The labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | EN ISO 18113‑2 Section 4.2 – Identification of kit components; US FDA 21 CFR Part 801 – Medical Device Labeling; IVDR Regulation (EU) 2017/746 Article 24 – Unique Device Identification (UDI) system | DP52-VEP-001 | 1 | Page 72-73
DP52 | PR.53 | The Product labelling must be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | US FDA 21 CFR Part 801 – Medical Device Labeling | DP52-VEP-001 | 1 | Page 73
DP52 | PR.54 | The Product labelling must be compliant with the relevant requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | IVDR Regulation (EU) 2017/746 – Regulation (EU) 2017/746 of the European Parliament and of the Council on in vitro diagnostic medical devices; ISO 15223‑1 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; ISO 18113 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) | DP52-VEP-001 | 1 | Page 73-75
DP52 | PR.55 | The Product labelling must be compliant with the requirements of a Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | DP52-VEP-001 | 1 | Page 75
DP52 | PR.56 | The Product labelling must be compliant with the relevant UK general IVD per the UK MDR 2002 labelling regulations. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | UK general IVD labelling regulations MDR 2002. | DP52-VEP-001 | 1 | Page 76
DP52 | PR.57 | The Product must only be compatible with BOND-III fully automated systems operating BOND Software that is compliant with regional regulatory requirements. | The Product can only be run on BOND software that is compliant with regional regulatory requirements. AND The Product Specifications should require the reagent container barcode to meet the DHFR 2119 BOND Barcode Specification. | LBS-MEL Software Verification; Product Specifications for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx (DP52-SPEC-001) | N/A | DP52-VEP-001 | 1 | Page 76
DP52 | PR.58 | The Product, when used within the HER2 BMA CDx Assay System, is expected to show greater agreement with reflex testing results, helping enable a more definitive determination of HER2 status based on the initial IHC test | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | DP52/DP61-VAP-002 Clinical Data Analysis Protocol for the HER2 BMA CDx Assay; LBS-CA-PTL-063.R02 Swift IHC Method Comparison Clinical Study Protocol | N/A | DP52-VEP-001 | 1 | Page 77
DP52 | PR.59 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | FFPE tissue suitable for use as positive control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | DP52-VEP-007 Precision Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 77-78
DP52 | PR.61 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | FDA guidance for IVD Device labelling; EU Labelling requirements for IVD; ISO 18113‑1 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements; ISO 18113‑2 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; EN ISO 15223‑1 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; EN ISO 14971 – Medical devices — Application of risk management to medical devices | DP52-VEP-001 | 1 | Page 78-80
DP52 | PR.62 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Current and Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | FDA guidance for IVD Device labelling; EU Labelling requirements for IVD; ISO 18113‑1 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements; ISO 18113‑2 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; EN ISO 15223‑1 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; EN ISO 14971 – Medical devices — Application of risk management to medical devices | DP52-VEP-001 | 1 | Page 81-83
DP52 | PR.63 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describes all required information. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | DP52‑VEP‑003 - Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx Protocol | FDA guidance for IVD Device labelling; EU Labelling requirements for IVD; ISO 18113‑1 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 1: Terms, definitions and general requirements; ISO 18113‑2 – In vitro diagnostic medical devices — Information supplied by the manufacturer (labelling) — Part 2: In vitro diagnostic reagents for professional use; EN ISO 15223‑1 – Medical devices — Symbols to be used with information to be supplied by the manufacturer — Part 1: General requirements; EN ISO 14971 – Medical devices — Application of risk management to medical devices | DP52-VEP-001 | 1 | Page 83-86
DP52 | PR.64 | The Product must be compatible with reagent‑grade alcohol, xylene, and xylene or toluene‑based mounting media. | All tissue cases must maintain acceptable staining when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA CDx Assay System, all slides must maintain acceptable staining when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA CDx Assay System. | DP52-VEP-008 - Robustness Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx | N/A | DP52-VEP-001 | 1 | Page 86-87
DP61 | PR.02 | The Product must contain cell cores representative of different HER2 expression levels (null HER2 expression, low HER2 expression and high HER2 expression). | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD | All verification studies generated using the HER2 BMA Cell Line Control will be incorporated into the Overarching Verification Reports for Project SWIFT (DP61‑VER‑001 and DP52‑VER‑001), where they will be assessed to verify that the cell cores did not contribute to any errors or anomalies. |  EU IVDR (Regulation (EU) 2017/746) — Annex I: General Safety and Performance Requirements  CLSI EP25 ED2:2023 — Evaluation of Stability of In Vitro Medical Laboratory Test Reagents  CLSI I/LA28 — Quality Assurance for Design Control (general robustness principles | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 11; Table 5, row 3 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.03 | The product, when used as part of the HER2 BMA CDx assay system, shall be capable of confirming correct execution of the assay when used in accordance with recommended assay parameters. | Deviations in the assay reagents, instrument and protocol that alter the HER2 expression status in patient samples should result in a failure of the CLC to meet the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑ROP‑001 Robustness Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑ROR‑001 Robustness Verification Report for the HER2 BMA Cell Line Control |  EU IVDR (Regulation (EU) 2017/746) — Annex I: General Safety and Performance Requirements  CLSI EP25 ED2:2023 — Evaluation of Stability of In Vitro Medical Laboratory Test Reagents  CLSI I/LA28 — Quality Assurance for Design Control (general robustness principles | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 11; Table 5, row 4 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.04 | The Product must be compatible for use with all components of the HER2 BMA CDx Assay System. | The Product, when stained using all components of the HER2 BMA Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | All verification studies generated using the HER2 BMA Cell Line Control will be incorporated into the Overarching Verification Reports for Project SWIFT (DP61‑VER‑001 and DP52‑VER‑001), where they will be assessed to verify that the HER2 BMA Cell Line Control is compatible within the Assay system. |  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 13485:2016 — Medical devices – Quality management systems  CLSI AUTO series (general guidance on automated IVD systems | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 11; Table 5, row 5 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.06 | The Product labelling must include images and describe how the product staining is interpreted. | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | Doc ID: DP61‑LABP‑001 Labelling Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑LABR‑001 Labelling Verification Report for the HER2 BMA Cell Line Control |  ISO 18113 1 / 2 — In vitro diagnostic medical devices — Information supplied by the manufacturer  EU IVDR (Regulation (EU) 2017/746) — Annex I, Sections 20–23  FDA 21 CFR Part 809 — In Vitro Diagnostic Products | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 11; Table 5, row 7 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.07 | The Product labelling must describe the expected staining patterns and characteristics of each cell cores. | HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to confirm they describes the expected staining patterns and characteristics of each cell line core. Sample Size: TBD | Doc ID: DP61‑LABP‑001 Labelling Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑LABR‑001 Labelling Verification Report for the HER2 BMA Cell Line Control |  ISO 18113 1 / 2 — In vitro diagnostic medical devices — Information supplied by the manufacturer  EU IVDR (Regulation (EU) 2017/746) — Annex I, Sections 20–23  FDA 21 CFR Part 809 — In Vitro Diagnostic Products | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 12; Table 5, row 8 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.08 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑STBP‑001 Stability Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑STBR‑001 Stability Verification Report for the HER2 BMA Cell Line Control | BS EN ISO 23640:2015 In vitro diagnostic medical devices. Evaluation of stability of in vitro diagnostic reagents CLSI EP25-ED2:2023 Evaluation of Stability of In Vitro Medical Laboratory Test Reagents EU IVDR (Regulation (EU) 2017/746) — Annex I: General Safety and Performance Requirements | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 12; Table 5, row 9 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.09 | The Product must generate Acceptable Staining without having to be subjected to any immediate preparation (e.g., slide baking or incubation) after being removed from the recommended storage conditions (2-8Â°C). | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | All verification studies generated using the HER2 BMA Cell Line Control will be incorporated into the Overarching Verification Reports for Project SWIFT (DP61‑VER‑001 and DP52‑VER‑001), where they will be assessed to verify that the HER2 BMA Cell Line Control is compatible with BOND-III fully automated system. |  BS EN ISO 23640:2015 — Evaluation of stability of IVD reagents  CLSI EP25 ED2:2023 — Evaluation of Stability of IVD Reagents  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 12; Table 5, row 10 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.10 | The Product must generate Acceptable Staining when used in conjunction with the recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols. | The Product, when stained using the HER2 BMA Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | All verification studies generated using the HER2 BMA Cell Line Control will be incorporated into the Overarching Verification Reports for Project SWIFT (DP61‑VER‑001 and DP52‑VER‑001), where they will be assessed to verify that the HER2 BMA Cell Line Control is compatible with BOND-III fully automated system. |  BS EN ISO 23640:2015 — Evaluation of stability of IVD reagents  CLSI EP25 ED2:2023 — Evaluation of Stability of IVD Reagents  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 13; Table 5, row 11 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.11 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life when stored in the immediate packaging (minimum of six months shelf life with a minimum two months safety margin). | Following a minimum of 14 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. The Product must also retains its ability to correctly detect procedural failures. Sample Size: TBD | Doc ID: DP61‑VEP‑005 Stability Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑005 Stability Verification Report for the HER2 BMA Cell Line Control |  BS EN ISO 23640:2015 — Evaluation of stability of IVD reagents  CLSI EP25 ED2:2023 — Evaluation of Stability of IVD Reagents  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 13; Table 5, row 12 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.12 | The Intermediate Product must be stable for a minimum of six months when stored at room temperature. | The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide after being stored at room temperature for at least six months. Sample Size: TBD | Doc ID: DP61‑VEP‑005 Stability Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑005 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP25 ED2:2023 — Evaluation of Stability of IVD Reagents  ISO 15189 principles (pre analytical specimen integrity)  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 13; Table 5, row 13 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.13 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | Three lots of the Product, manufactured from separate multiblocks will be stained using a single lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by one trained pathologist, with the lot of the Product as the only variable. Staining results must align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑VEP‑003 Precision Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑003 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP05 A3 — Evaluation of Precision Performance of Quantitative Measurement Methods  CLSI EP12 A2 — User Protocol for Evaluation of Qualitative Test Performance  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 5725 1 — Accuracy (trueness and precision) of measurement methods | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 14; Table 5, row 14 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.14 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining throughout the same lots of The Product (intra-lot repeatability). | Sections taken from the beginning, middle, and end of a single Cell Line Control multiblock will be stained using one lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. A trained pathologist will evaluate the slides, with the section position within the multiblock as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑VEP‑003 Precision Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑003 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP05 A3 — Evaluation of Precision Performance of Quantitative Measurement Methods  CLSI EP12 A2 — User Protocol for Evaluation of Qualitative Test Performance  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 5725 1 — Accuracy (trueness and precision) of measurement methods | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 14; Table 5, row 15 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.15 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | Cell Line Control slides from the same lot will be stained on five non-consecutive days (spanning a minimum period of 20 days). A single trained pathologist will score the slides, with the staining day being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑VEP‑003 Precision Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑003 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP05 A3 — Evaluation of Precision Performance of Quantitative Measurement Methods  CLSI EP12 A2 — User Protocol for Evaluation of Qualitative Test Performance  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 5725 1 — Accuracy (trueness and precision) of measurement methods | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 14; Table 5, row 16 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.16 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | Cell Line Control slides from the same lot will be stained using three different BOND-III fully automated systems. A single trained pathologist will score the slides, with the BOND-III instrument being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑VEP‑003 Precision Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑003 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP05 A3 — Evaluation of Precision Performance of Quantitative Measurement Methods  CLSI EP12 A2 — User Protocol for Evaluation of Qualitative Test Performance  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 5725 1 — Accuracy (trueness and precision) of measurement methods | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 15; Table 5, row 17 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.17 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | Cell Line Control slides from the same lot will be stained on different BOND Runs/Slide Staining Assemblies (SSAs) within the same BOND Run Batch. A single trained pathologist will score the slides, with the BOND Run/SSA as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑VEP‑003 Precision Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑003 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP05 A3 — Evaluation of Precision Performance of Quantitative Measurement Methods  CLSI EP12 A2 — User Protocol for Evaluation of Qualitative Test Performance  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 5725 1 — Accuracy (trueness and precision) of measurement methods | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 15; Table 5, row 18 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.18 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Polymer Refine Detection (DS9800) system. The slides will be scored by a single trained pathologist, with the lot of the DS9800 as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑VEP‑003 Precision Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑003 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP05 A3 — Evaluation of Precision Performance of Quantitative Measurement Methods  CLSI EP12 A2 — User Protocol for Evaluation of Qualitative Test Performance  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 5725 1 — Accuracy (trueness and precision) of measurement methods | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 15; Table 5, row 19 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.19 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx are varied (inter-antibody lot repeatability). | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by a single trained pathologist, with the lot of the antibody as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | Doc ID: DP61‑VEP‑003 Precision Verification Plan for the HER2 BMA Cell Line Control Doc ID: DP61‑VER‑003 Stability Verification Report for the HER2 BMA Cell Line Control |  CLSI EP05 A3 — Evaluation of Precision Performance of Quantitative Measurement Methods  CLSI EP12 A2 — User Protocol for Evaluation of Qualitative Test Performance  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 5725 1 — Accuracy (trueness and precision) of measurement methods | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 16; Table 5, row 20 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.20 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III temperature specifications. | Cell Line Control slides from BOND runs performed at environmental temperatures between 18-26°C must stain as expected - Core A (BT474-M3) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  |  IEC 61010 1 — Safety requirements for laboratory equipment (environmental conditions)  EU IVDR (Regulation (EU) 2017/746) — Annex I  Manufacturer’s instrument specifications (BOND III) | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 16; Table 5, row 21 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.21 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III humidity specifications. | Cell Line Control slides from BOND runs performed at environmental humidities between 30-80% Relative Humidity must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD |  |  IEC 61010 1 — Safety requirements for laboratory equipment (environmental conditions)  EU IVDR (Regulation (EU) 2017/746) — Annex I  Manufacturer’s instrument specifications (BOND III) | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 16; Table 5, row 22 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.22 | The placement of the cell line cores within The Product must be consistently orientated relative to each other. | HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA Assay Interpretation Guide. Sample Size: TBD |  |  ISO 15189 — Quality and competence in medical laboratories  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 16; Table 5, row 23 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.23 | The Product must have labelling that provides all required hazard and safety information and associated symbols. | Users will review the Product slide mailer label, hazard notification slip and IFU and complete a checklist to confirm that the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents) is present. Sample Size: TBD |  |  CLP Regulation (EC) No 1272/2008  REACH Regulation (EC) No 1907/2006  GHS (Globally Harmonized System of Classification and Labelling of Chemicals)  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 17; Table 5, row 24 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.24 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. Sample Size: TBD |  |  CLP Regulation (EC) No 1272/2008  REACH Regulation (EC) No 1907/2006  GHS (Globally Harmonized System of Classification and Labelling of Chemicals)  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 17; Table 5, row 25 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.25 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD |  |  ISO 23640:2015 — Stability and transport simulation for IVDs  ISTA 2A / 3A — Package transport simulation  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 17; Table 5, row 26 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.26 | The outer packaging must be robust enough to ensure the containment of the Product if damaged or broken during routine shipping and storage. | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD |  |  ISO 23640:2015 — Stability and transport simulation for IVDs  ISTA 2A / 3A — Package transport simulation  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 18; Table 5, row 27 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.27 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD |  |  ISO 23640:2015 — Stability and transport simulation for IVDs  ISTA 2A / 3A — Package transport simulation  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 18; Table 5, row 28 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.28 | The on-slide printed text must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired. | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. Sample Size: TBD |  |  ISO 18113 1 / 2 — IVD labelling requirements  ISO 15223 1 — Symbols to be used with medical device labels  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 18; Table 5, row 29 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.29 | The Immediate Packaging label must be capable of withstanding shipping and storage within the recommended temperature range and normal use without becoming detached. | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remain fully adhered and undamaged. Sample Size: TBD |  |  ISO 18113 1 / 2 — IVD labelling requirements  ISO 15223 1 — Symbols to be used with medical device labels  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 18; Table 5, row 30 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.30 | The symbols and text included on The Product labelling must be human legible. | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. Sample Size: TBD |  |  ISO 18113 1 / 2 — IVD labelling requirements  ISO 15223 1 — Symbols to be used with medical device labels  EU IVDR (Regulation (EU) 2017/746) — Annex I | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 19; Table 5, row 31 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.31 | The Product shall include compliance with UDI specifications. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. Sample Size: TBD |  |  EU IVDR (Regulation (EU) 2017/746) — Article 24  FDA 21 CFR Part 801 — UDI requirements  ISO 18113 2 — IVD labelling | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 19; Table 5, row 32 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.32 | The Product labelling must be created to be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. Sample Size: TBD |  |  US: FDA 21 CFR Part 801 & 809  EU: IVDR (EU) 2017/746  Canada: SOR/98 282  UK: UK MDR 2002  ISO 18113 series | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 19; Table 5, row 33 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.33 | The Product labelling must be created to be compliant with the relevant requirements from IVDR EU/2017/746. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. Sample Size: TBD |  |  US: FDA 21 CFR Part 801 & 809  EU: IVDR (EU) 2017/746  Canada: SOR/98 282  UK: UK MDR 2002  ISO 18113 series | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 20; Table 5, row 34 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.34 | The Product labelling must be created to be compliant with the relevant Canadian IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. Sample Size: TBD |  |  US: FDA 21 CFR Part 801 & 809  EU: IVDR (EU) 2017/746  Canada: SOR/98 282  UK: UK MDR 2002  ISO 18113 series | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 20; Table 5, row 35 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.35 | The Product labelling must be created to comply with the relevant UK general IVD as per the UK MDR 2002 labelling regulations. | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. Sample Size: TBD |  |  US: FDA 21 CFR Part 801 & 809  EU: IVDR (EU) 2017/746  Canada: SOR/98 282  UK: UK MDR 2002  ISO 18113 series | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 20; Table 5, row 36 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.36 | The Product, when used within the HER2 BMA CDx Assay System, must generate acceptable staining when used within the BOND-III altitude specifications. | Cell Line Control slides from BOND runs performed at altitudes up to and including 1600 meters above sea level must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBC |  |  IEC 61010 1 — Safety requirements for laboratory equipment (environmental conditions)  EU IVDR (Regulation (EU) 2017/746) — Annex I  Manufacturer’s instrument specifications (BOND III) | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 21; Table 5, row 37 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.37 | The Product Instruction for Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD |  |  ISO 14971:2019 — Application of risk management to medical devices  EU IVDR (Regulation (EU) 2017/746) — Annex I  ISO 18113 series | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 21; Table 5, row 38 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.
DP61 | PR.38 | The Product must be compatible with reagent grade alcohol, xylene and toluene based mounting media. | All Cell Line Controls must maintain acceptable staining and cell morphology when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA Assay System. All Cell Line Controls must maintain acceptable staining and cell morphology when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA Assay System. Sample Size: TBD |  |  | DP61-PR-001 / DP61-VEP-001 | 05 / [NEEDS INPUT] | PR page 21; Table 5, row 39 | Requirement text from signed DP61-PR-001 Rev 05; acceptance criteria from supplied DP61 overarching plan.

## PR-Risk Links (hidden)
ProductFamily | ProductRequirementID | RiskSourceDocument | RiskID | RiskExistsInWorkbook | RiskWorkbookExcelRow | RiskRawIdentifier | RiskIdentifierStatus | VerificationStatus | SourceDocument | SourceRevision | SourcePage
DP52 | PR.01 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 11
DP52 | PR.01 | LBN-dRMW-016 | dRMW-risk-002 | Yes | 9 | dRMW-risk-002 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 11
DP52 | PR.01 | LBN-dRMW-016 | dRMW-risk-003 | Yes | 10 | dRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 11
DP52 | PR.01 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 11
DP52 | PR.01 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 11
DP52 | PR.01 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | dRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 11
DP52 | PR.01 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 11
DP52 | PR.02 | LBN-dRMW-016 | dRMW-risk-023 | Yes | 20 | dRMW-risk-023 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | dRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | dRMW-risk-025 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | dRMW-risk-034 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-015 | Yes | 16 | euRMW-risk-015 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-017 | Yes | 17 | euRMW-risk-017 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-048 | Yes | 35 | euRMW-risk-048 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-049 | Yes | 36 | euRMW-risk-049 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-050 | Yes | 37 | euRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-052 | Yes | 38 | euRMW-risk-052 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-058 | Yes | 41 | euRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-064 | Yes | 42 | euRMW-risk-064 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-066 | Yes | 43 | euRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.02 | LBN-euRMW-020 | euRMW-risk-080 | Yes | 49 | euRMW-risk-080 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.03 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.03 | LBN-dRMW-016 | dRMW-risk-002 | Yes | 9 | dRMW-risk-002 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.03 | LBN-dRMW-016 | dRMW-risk-003 | Yes | 10 | dRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.03 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.03 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.03 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | dRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.03 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 12
DP52 | PR.04 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | dRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-dRMW-016 | dRMW-risk-028 | Yes | 23 | dRMW-risk-028 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-dRMW-016 | dRMW-risk-039 | Yes | 27 | dRMW-risk-039 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-dRMW-016 | dRMW-risk-040 | Yes | 28 | dRMW-risk-040 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | euRMW-risk-013 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-euRMW-020 | euRMW-risk-019 | Yes | 18 | euRMW-risk-019 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-euRMW-020 | euRMW-risk-019.5 | Yes | 19 | euRMW-risk-019.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-euRMW-020 | euRMW-risk-020 | Yes | 20 | euRMW-risk-020 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-euRMW-020 | euRMW-risk-021 | Yes | 21 | euRMW-risk-021 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-euRMW-020 | euRMW-risk-022.1 | Yes | 22 | euRMW-risk-022.1 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.04 | LBN-euRMW-020 | euRMW-risk-023 | Yes | 23 | euRMW-risk-023 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-003 | Yes | 10 | dRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | dRMW-risk-020.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | dRMW-risk-025 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-029 | Yes | 24 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | dRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | dRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | dRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-050 | Yes | 32 | dRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-053 | Yes | 33 | dRMW-risk-053 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | dRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | dRMW-risk-059 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | dRMW-risk-060 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | dRMW-risk-062 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | dRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | dRMW-risk-069 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-euRMW-020 | euRMW-risk-011.5 | Yes | 14 | euRMW-risk-011.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | euRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.05 | LBN-euRMW-020 | euRMW-risk-072 | Yes | 46 | euRMW-risk-072 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 13
DP52 | PR.06 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-dRMW-016 | dRMW-risk-002 | Yes | 9 | dRMW-risk-002 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-dRMW-016 | dRMW-risk-003 | Yes | 10 | dRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | dRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-euRMW-020 | euRMW-risk-001 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | euRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-euRMW-020 | euRMW-risk-005 | Yes | 9 | euRMW-risk-005 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-euRMW-020 | euRMW-risk-009 | Yes | 10 | euRMW-risk-009 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-euRMW-020 | euRMW-risk-011 | Yes | 11 | euRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.06 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | euRMW-risk-013 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-euRMW-020 | euRMW-risk-001 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | euRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-euRMW-020 | euRMW-risk-009 | Yes | 10 | euRMW-risk-009 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-euRMW-020 | euRMW-risk-011 | Yes | 11 | euRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.07 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | euRMW-risk-013 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.08 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.08 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.08 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.08 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 14
DP52 | PR.08 | LBN-euRMW-020 | euRMW-risk-001 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 14
DP52 | PR.10 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | dRMW-risk-060 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.11 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | dRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.11 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | dRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.11 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | dRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.11 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | dRMW-risk-069 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.11 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | euRMW-risk-091 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.12 | LBN-dRMW-016 | dRMW-risk-023 | Yes | 20 | dRMW-risk-023 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.12 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | dRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.12 | LBN-euRMW-020 | euRMW-risk-068 | Yes | 44 | euRMW-risk-068 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 15
DP52 | PR.13 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | dRMW-risk-062 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.13 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | dRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.13 | LBN-euRMW-020 | euRMW-risk-092 | Yes | 54 | euRMW-risk-092 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.14 | LBN-dRMW-016 | dRMW-risk-019 | Yes | 15 | dRMW-risk-019 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.14 | LBN-dRMW-016 | dRMW-risk-020 | Yes | 16 | dRMW-risk-020 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.14 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | dRMW-risk-020.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.14 | LBN-euRMW-020 | euRMW-risk-064 | Yes | 42 | euRMW-risk-064 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.15 | LBN-dRMW-016 | dRMW-risk-019 | Yes | 15 | dRMW-risk-019 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.15 | LBN-dRMW-016 | dRMW-risk-020 | Yes | 16 | dRMW-risk-020 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.15 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | dRMW-risk-020.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.15 | LBN-dRMW-016 | dRMW-risk-021 | Yes | 18 | dRMW-risk-021 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.15 | LBN-euRMW-020 | euRMW-risk-064 | Yes | 42 | euRMW-risk-064 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 16
DP52 | PR.16 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | dRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 17
DP52 | PR.16 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | dRMW-risk-059 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 17
DP52 | PR.16 | LBN-euRMW-020 | euRMW-risk-042 | Yes | 30 | euRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 17
DP52 | PR.16 | LBN-euRMW-020 | euRMW-risk-043 | Yes | 31 | euRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 17
DP52 | PR.17 | LBN-euRMW-020 | euRMW-risk-055 | Yes | 40 | euRMW-risk-055 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 17
DP52 | PR.18 | LBN-euRMW-020 | euRMW-risk-055 | Yes | 40 | euRMW-risk-055 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 17
DP52 | PR.18 | LBN-euRMW-020 | euRMW-risk-093 | Yes | 55 | euRMW-risk-093 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 17
DP52 | PR.19 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | dRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.19 | LBN-dRMW-016 | dRMW-risk-070 | Yes | 43 | dRMW-risk-070 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.19 | LBN-dRMW-016 | dRMW-risk-070.5 | Yes | 44 | dRMW-risk-070.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.20 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | dRMW-risk-034 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.20 | LBN-euRMW-020 | euRMW-risk-066 | Yes | 43 | euRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.20 | LBN-euRMW-020 | euRMW-risk-078 | Yes | 48 | euRMW-risk-078 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.21 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | dMRW-risk-086 | [NEEDS INPUT] Source identifier says dMRW; canonicalized to dRMW for linkage | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.21 | LBN-euRMW-020 | euRMW-risk-001 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 18
DP52 | PR.21 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | euRMW-risk-089 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.21 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | euRMW-risk-098 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 18
DP52 | PR.22 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | dRMW-risk-034 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 19
DP52 | PR.22 | LBN-euRMW-020 | euRMW-risk-078 | Yes | 48 | euRMW-risk-078 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 19
DP52 | PR.22 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | euRMW-risk-089 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 19
DP52 | PR.22 | LBN-euRMW-020 | euRMW-risk-097 | Yes | 57 | euRMW-risk-097 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 19
DP52 | PR.23 | LBN-dRMW-016 | dRMW-risk-039 | Yes | 27 | dRMW-risk-039 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 19
DP52 | PR.24 | LBN-dRMW-016 | dRMW-risk-040 | Yes | 28 | dRMW-risk-040 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 19
DP52 | PR.25 | LBN-dRMW-016 | dRMW-risk-029 | Yes | 24 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 20
DP52 | PR.26 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 20
DP52 | PR.27 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | dRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 20
DP52 | PR.28 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | dRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 21
DP52 | PR.30 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | dRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 21
DP52 | PR.31 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | dRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 22
DP52 | PR.35 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | dMRW-risk-086 | [NEEDS INPUT] Source identifier says dMRW; canonicalized to dRMW for linkage | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.35 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | euRMW-risk-098 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.36 | LBN-dRMW-016 | dRMW-risk-088 | Yes | 59 | dMRW-risk-088 dRMW-risk-088 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.36 | LBN-euRMW-020 | euRMW-risk-083 | Yes | 50 | euRMW-risk-083 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.37 | LBN-dRMW-016 | dRMW-risk-089 | Yes | 60 | dMRW-risk-089 dRMW-risk-089 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.38 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | dMRW-risk-086 | [NEEDS INPUT] Source identifier says dMRW; canonicalized to dRMW for linkage | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.38 | LBN-dRMW-016 | dRMW-risk-088 | Yes | 59 | dMRW-risk-088 dRMW-risk-088 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.38 | LBN-dRMW-016 | dRMW-risk-089 | Yes | 60 | dMRW-risk-089 dRMW-risk-089 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 23
DP52 | PR.39 | LBN-dRMW-016 | dRMW-risk-056 | Yes | 34 | dRMW-risk-056 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 24
DP52 | PR.39 | LBN-dRMW-016 | dRMW-risk-057 | Yes | 35 | dRMW-risk-057 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 24
DP52 | PR.41 | LBN-euRMW-020 | euRMW-risk-093 | Yes | 55 | euRMW-risk-093 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 24
DP52 | PR.42 | LBN-euRMW-020 | euRMW-risk-096 | Yes | 56 | euRMW-risk-096 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 24
DP52 | PR.43 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | dRMW-risk-071 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.43 | LBN-dRMW-016 | dRMW-risk-072 | Yes | 46 | dRMW-risk-072 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.43 | LBN-dRMW-016 | dRMW-risk-073 | Yes | 47 | dRMW-risk-073 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.43 | LBN-dRMW-016 | dRMW-risk-074 | Yes | 48 | dRMW-risk-074 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.43 | LBN-dRMW-016 | dRMW-risk-075 | Yes | 49 | dRMW-risk-075 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.43 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | euRMW-risk-044 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.43 | LBN-euRMW-020 | euRMW-risk-046 | Yes | 33 | euRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.43 | LBN-euRMW-020 | euRMW-risk-047 | Yes | 34 | euRMW-risk-047 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 25
DP52 | PR.44 | LBN-dRMW-016 | dRMW-risk-078 | Yes | 51 | dRMW-risk-078 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 26
DP52 | PR.44 | LBN-dRMW-016 | dRMW-risk-081 | Yes | 53 | dRMW-risk-081 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 26
DP52 | PR.45 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | dRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 26
DP52 | PR.45 | LBN-euRMW-020 | euRMW-risk-042 | Yes | 30 | euRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 26
DP52 | PR.46 | LBN-dRMW-016 | dRMW-risk-068 | Yes | 41 | dRMW-risk-068 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.46 | LBN-euRMW-020 | euRMW-risk-033.5 | Yes | 27 | euRMW-risk-033.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.46 | LBN-euRMW-020 | euRMW-risk-039 | Yes | 29 | euRMW-risk-039 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.46 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | euRMW-risk-044 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.46 | LBN-euRMW-020 | euRMW-risk-046 | Yes | 33 | euRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.46 | LBN-euRMW-020 | euRMW-risk-047 | Yes | 34 | euRMW-risk-047 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.46 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | euRMW-risk-091 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.47 | LBN-dRMW-016 | dRMW-risk-084 | Yes | 55 | dRMW-risk-084 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.47 | LBN-dRMW-016 | dRMW-risk-085 | Yes | 56 | dRMW-risk-085 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.47 | LBN-euRMW-020 | euRMW-risk-039 | Yes | 29 | euRMW-risk-039 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 27
DP52 | PR.48 | LBN-dRMW-016 | dRMW-risk-084 | Yes | 55 | dRMW-risk-084 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.48 | LBN-dRMW-016 | dRMW-risk-085 | Yes | 56 | dRMW-risk-085 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.48 | LBN-dRMW-016 | dRMW-risk-090 | Yes | 61 | dMRW-risk-090 dRMW-risk-090 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.48 | LBN-dRMW-016 | dRMW-risk-091 | Yes | 62 | dMRW-risk-091 dRMW-risk-091 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.48 | LBN-euRMW-020 | euRMW-risk-033.5 | Yes | 27 | euRMW-risk-033.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.49 | LBN-dRMW-016 | dRMW-risk-092 | Yes | 63 | dMRW-risk-092 dRMW-risk-092 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.49 | LBN-euRMW-020 | euRMW-risk-032.5 | Yes | 26 | euRMW-risk-032.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.50 | LBN-dRMW-016 | dRMW-risk-080 | Yes | 52 | dRMW-risk-080 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 28
DP52 | PR.51 | LBN-dRMW-016 | dRMW-risk-078 | Yes | 51 | dRMW-risk-078 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 29
DP52 | PR.51 | LBN-dRMW-016 | dRMW-risk-081 | Yes | 53 | dRMW-risk-081 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 29
DP52 | PR.51 | LBN-dRMW-016 | dRMW-risk-082 | Yes | 54 | dRMW-risk-082 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 29
DP52 | PR.53 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | dRMW-risk-071 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 29
DP52 | PR.54 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | dRMW-risk-071 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 30
DP52 | PR.55 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | dRMW-risk-071 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 30
DP52 | PR.56 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | dRMW-risk-071 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 31
DP52 | PR.57 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | dRMW-risk-071 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 31
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | dRMW-risk-020.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | dRMW-risk-025 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-029 | Yes | 24 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | dRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | dRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | dRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-050 | Yes | 32 | dRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-053 | Yes | 33 | dRMW-risk-053 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | dRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | dRMW-risk-059 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | dRMW-risk-060 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | dRMW-risk-062 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | dRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | dRMW-risk-069 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-001 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | euRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-011.5 | Yes | 14 | euRMW-risk-011.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-021 | Yes | 21 | euRMW-risk-021 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-022.1 | Yes | 22 | euRMW-risk-022.1 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-042 | Yes | 30 | euRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-075 | Yes | 47 | euRMW-risk-075 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-080 | Yes | 49 | euRMW-risk-080 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-083 | Yes | 50 | euRMW-risk-083 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | euRMW-risk-089 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | euRMW-risk-091 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-097 | Yes | 57 | euRMW-risk-097 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.59 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | euRMW-risk-098 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | dRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-dRMW-016 | dRMW-risk-028 | Yes | 23 | dRMW-risk-028 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-dRMW-016 | dRMW-risk-039 | Yes | 27 | dRMW-risk-039 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-dRMW-016 | dRMW-risk-040 | Yes | 28 | dRMW-risk-040 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-euRMW-020 | euRMW-risk-019 | Yes | 18 | euRMW-risk-019 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-euRMW-020 | euRMW-risk-019.5 | Yes | 19 | euRMW-risk-019.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-euRMW-020 | euRMW-risk-020 | Yes | 20 | euRMW-risk-020 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-euRMW-020 | euRMW-risk-021 | Yes | 21 | euRMW-risk-021 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-euRMW-020 | euRMW-risk-022.1 | Yes | 22 | euRMW-risk-022.1 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-euRMW-020 | euRMW-risk-023 | Yes | 23 | euRMW-risk-023 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.61 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | euRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 32
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | dRMW-risk-016 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-019 | Yes | 15 | dRMW-risk-019 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | dRMW-risk-020.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-023 | Yes | 20 | dRMW-risk-023 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | dRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | dRMW-risk-025 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-028 | Yes | 23 | dRMW-risk-028 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-029 | Yes | 24 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | dRMW-risk-034 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | dRMW-risk-042 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | dRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | dRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-050 | Yes | 32 | dRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-053 | Yes | 33 | dRMW-risk-053 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | dRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | dRMW-risk-059 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | dRMW-risk-060 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | dRMW-risk-062 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | dRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-068 | Yes | 41 | dRMW-risk-068 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | dRMW-risk-069 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-084 | Yes | 55 | dRMW-risk-084 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | dMRW-risk-086 | [NEEDS INPUT] Source identifier says dMRW; canonicalized to dRMW for linkage | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-087 | Yes | 58 | dMRW-risk-087 dRMW-risk-087 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-088 | Yes | 59 | dMRW-risk-088 dRMW-risk-088 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-dRMW-016 | dRMW-risk-089 | Yes | 60 | dMRW-risk-089 dRMW-risk-089 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-001 | No |  |  |  | [NEEDS INPUT] PR citation not found in supplied workbook | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | euRMW-risk-003 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-005 | Yes | 9 | euRMW-risk-005 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-009 | Yes | 10 | euRMW-risk-009 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-011 | Yes | 11 | euRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-011.5 | Yes | 14 | euRMW-risk-011.5 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | euRMW-risk-013 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-015 | Yes | 16 | euRMW-risk-015 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-017 | Yes | 17 | euRMW-risk-017 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | euRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | euRMW-risk-044 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-046 | Yes | 33 | euRMW-risk-046 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-047 | Yes | 34 | euRMW-risk-047 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-048 | Yes | 35 | euRMW-risk-048 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-049 | Yes | 36 | euRMW-risk-049 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-050 | Yes | 37 | euRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-052 | Yes | 38 | euRMW-risk-052 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-053 | Yes | 39 | euRMW-risk-053 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-055 | Yes | 40 | euRMW-risk-055 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-058 | Yes | 41 | euRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-064 | Yes | 42 | euRMW-risk-064 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-066 | Yes | 43 | euRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-068 | Yes | 44 | euRMW-risk-068 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-070 | Yes | 45 | euRMW-risk-070 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-072 | Yes | 46 | euRMW-risk-072 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-075 | Yes | 47 | euRMW-risk-075 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-078 | Yes | 48 | euRMW-risk-078 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-080 | Yes | 49 | euRMW-risk-080 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-083 | Yes | 50 | euRMW-risk-083 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-086 | Yes | 51 | euRMW-risk-086 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | euRMW-risk-089 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | euRMW-risk-091 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-092 | Yes | 54 | euRMW-risk-092 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-093 | Yes | 55 | euRMW-risk-093 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-096 | Yes | 56 | euRMW-risk-096 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-097 | Yes | 57 | euRMW-risk-097 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.62 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | euRMW-risk-098 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 33
DP52 | PR.63 | LBN-dRMW-016 | dRMW-risk-068 | Yes | 41 | dRMW-risk-068 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 34
DP52 | PR.63 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | dRMW-risk-071 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 34
DP52 | PR.63 | LBN-dRMW-016 | dRMW-risk-081 | Yes | 53 | dRMW-risk-081 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 34
DP52 | PR.63 | LBN-dRMW-016 | dRMW-risk-084 | Yes | 55 | dRMW-risk-084 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 34
DP52 | PR.63 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | euRMW-risk-044 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 34
DP52 | PR.64 | LBN-dRMW-016 | dRMW-risk-087 | Yes | 58 | dMRW-risk-087 dRMW-risk-087 | Source cell contains identifier variants; canonical valid identifier selected | Verified against supplied workbook row | DP52-PR-001 | 6 | 34
DP52 | PR.64 | LBN-euRMW-020 | euRMW-risk-053 | Yes | 39 | euRMW-risk-053 | Exact source identifier | Verified against supplied workbook row | DP52-PR-001 | 6 | 34
DP61 | PR.02 | LBN-dRMW-017 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.02 | LBN-dRMW-017 | dRMW-risk-009 | Yes | 12 | dRMW-risk-009 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.02 | LBN-dRMW-017 | dRMW-risk-010 | Yes | 13 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.03 | LBN-dRMW-017 | dRMW-risk-010 | Yes | 13 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.03 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | euRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.04 | LBN-dRMW-017 | dRMW-risk-009 | Yes | 12 | dRMW-risk-009 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.06 | LBN-dRMW-017 | dRMW-risk-013 | Yes | 15 | dRMW-risk-013 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.06 | LBN-dRMW-017 | dRMW-risk-040 | Yes | 32 | dRMW-risk-040 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.06 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | euRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 11
DP61 | PR.07 | LBN-dRMW-017 | dRMW-risk-013 | Yes | 15 | dRMW-risk-013 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 12
DP61 | PR.07 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | euRMW-risk-024 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 12
DP61 | PR.08 | LBN-dRMW-017 | dRMW-risk-018 | Yes | 17 | dRMW-risk-018 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 12
DP61 | PR.09 | LBN-dRMW-017 | dRMW-risk-001 | Yes | 8 | dRMW-risk-001 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 12
DP61 | PR.09 | LBN-dRMW-017 | dRMW-risk-004 | Yes | 9 | dRMW-risk-004 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 12
DP61 | PR.11 | LBN-dRMW-017 | dRMW-risk-017 | Yes | 16 | dRMW-risk-017 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 13
DP61 | PR.11 | LBN-euRMW-020 | euRMW-risk-043 | Yes | 31 | euRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 13
DP61 | PR.12 | LBN-dRMW-017 | dRMW-risk-039 | Yes | 31 | dRMW-risk-039 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 13
DP61 | PR.13 | LBN-dRMW-017 | dRMW-risk-021 | Yes | 18 | dRMW-risk-021 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 14
DP61 | PR.13 | LBN-dRMW-017 | dRMW-risk-039 | Yes | 31 | dRMW-risk-039 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 14
DP61 | PR.15 | LBN-dRMW-017 | dRMW-risk-022 | Yes | 21 | dRMW-risk-022 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 14
DP61 | PR.16 | LBN-dRMW-017 | dRMW-risk-022 | Yes | 21 | dRMW-risk-022 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 15
DP61 | PR.18 | LBN-dRMW-017 | dRMW-risk-021.2 | Yes | 20 | dRMW-risk-021.2 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 15
DP61 | PR.19 | LBN-dRMW-017 | dRMW-risk-021.1 | Yes | 19 | dRMW-risk-021.1 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 16
DP61 | PR.20 | LBN-dRMW-017 | dRMW-risk-047 | Yes | 39 | dRMW-risk-047 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 16
DP61 | PR.21 | LBN-dRMW-017 | dRMW-risk-048 | Yes | 40 | dRMW-risk-048 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 16
DP61 | PR.22 | LBN-dRMW-017 | dRMW-risk-006 | Yes | 10 | dRMW-risk-006 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 16
DP61 | PR.23 | LBN-dRMW-017 | dRMW-risk-028 | Yes | 22 | dRMW-risk-028 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 17
DP61 | PR.24 | LBN-dRMW-017 | dRMW-risk-028 | Yes | 22 | dRMW-risk-028 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 17
DP61 | PR.25 | LBN-dRMW-017 | dRMW-risk-032 | Yes | 26 | dRMW-risk-032 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 17
DP61 | PR.25 | LBN-dRMW-017 | dRMW-risk-032.5 | Yes | 27 | dRMW-risk-032.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 17
DP61 | PR.25 | LBN-dRMW-017 | dRMW-risk-033 | Yes | 28 | dRMW-risk-033 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 17
DP61 | PR.25 | LBN-dRMW-017 | dRMW-risk-033.5 | Yes | 29 | dRMW-risk-033.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 17
DP61 | PR.26 | LBN-dRMW-017 | dRMW-risk-032 | Yes | 26 | dRMW-risk-032 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.26 | LBN-dRMW-017 | dRMW-risk-032.5 | Yes | 27 | dRMW-risk-032.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.26 | LBN-dRMW-017 | dRMW-risk-033 | Yes | 28 | dRMW-risk-033 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.26 | LBN-dRMW-017 | dRMW-risk-033.5 | Yes | 29 | dRMW-risk-033.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.27 | LBN-dRMW-017 | dRMW-risk-032.5 | Yes | 27 | dRMW-risk-032.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.27 | LBN-dRMW-017 | dRMW-risk-050 | Yes | 44 | dRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.27 | LBN-dRMW-017 | dRMW-risk-051 | Yes | 45 | dRMW-risk-051 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.28 | LBN-dRMW-017 | dRMW-risk-031 | Yes | 25 | dRMW-risk-031 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.28 | LBN-dRMW-017 | dRMW-risk-041 | Yes | 33 | dRMW-risk-041 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.29 | LBN-dRMW-017 | dRMW-risk-043 | Yes | 35 | dRMW-risk-043 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.29 | LBN-dRMW-017 | dRMW-risk-044 | Yes | 36 | dRMW-risk-044 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 18
DP61 | PR.30 | LBN-dRMW-017 | dRMW-risk-041 | Yes | 33 | dRMW-risk-041 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 19
DP61 | PR.32 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 19
DP61 | PR.33 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 20
DP61 | PR.34 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 20
DP61 | PR.35 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | dRMW-risk-029 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 20
DP61 | PR.36 | LBN-dRMW-017 | dRMW-risk-049 | Yes | 41 | dRMW-risk-049 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-004 | Yes | 9 | dRMW-risk-004 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-006 | Yes | 10 | dRMW-risk-006 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-007 | Yes | 11 | dRMW-risk-007 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-010 | Yes | 13 | dRMW-risk-010 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-011 | Yes | 14 | dRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-013 | Yes | 15 | dRMW-risk-013 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-028 | Yes | 22 | dRMW-risk-028 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-033 | Yes | 28 | dRMW-risk-033 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-033.5 | Yes | 29 | dRMW-risk-033.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-049 | Yes | 41 | dRMW-risk-049 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-049.01 | Yes | 42 | dRMW-risk-049.01 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-049.1 | Yes | 43 | dRMW-risk-049.1 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-050 | Yes | 44 | dRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-dRMW-017 | dRMW-risk-052 | Yes | 46 | dRMW-risk-052 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-011.5 | Yes | 14 | euRMW-risk-011.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-024.5 | Yes | 25 | euRMW-risk-024.5 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-047 | Yes | 34 | euRMW-risk-047 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-048 | Yes | 35 | euRMW-risk-048 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-049 | Yes | 36 | euRMW-risk-049 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-050 | Yes | 37 | euRMW-risk-050 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-052 | Yes | 38 | euRMW-risk-052 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-053 | Yes | 39 | euRMW-risk-053 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-058 | Yes | 41 | euRMW-risk-058 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-066 | Yes | 43 | euRMW-risk-066 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-068 | Yes | 44 | euRMW-risk-068 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-070 | Yes | 45 | euRMW-risk-070 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-072 | Yes | 46 | euRMW-risk-072 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-075 | Yes | 47 | euRMW-risk-075 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-078 | Yes | 48 | euRMW-risk-078 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-080 | Yes | 49 | euRMW-risk-080 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-083 | Yes | 50 | euRMW-risk-083 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-086 | Yes | 51 | euRMW-risk-086 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | euRMW-risk-089 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | euRMW-risk-091 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-096 | Yes | 56 | euRMW-risk-096 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-097 | Yes | 57 | euRMW-risk-097 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.37 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | euRMW-risk-098 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21
DP61 | PR.38 | LBN-dRMW-017 | dRMW-risk-011 | Yes | 14 | dRMW-risk-011 | Exact source identifier | Verified against supplied workbook row | DP61-PR-001 | 5 | 21

## Protocol Coverage (hidden)
ProductFamily | ProtocolID | ProductRequirementID | RiskReferenceText | ProductRequirementTextInProtocol | AcceptanceCriteriaInProtocol | ParsedRiskLinkCount | SourceTable | SourceTableRow | SourceFile
DP52 | DP52.DP61-VAP-002 | PR.58 | N/A – Marketing Requirement | The Product, when used within the HER2 BMA CDx Assay System, is expected to show greater agreement with reflex testing results, helping enable a more definitive determination of HER2 status based on the initial IHC test. | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | 0 | 5 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52 | DP52-VEP-003 | PR.04 | LBN-dRMW-016 Risk Lines 010, 011, 016, 028, 039 and 040 LBN-euRMW-020 Risk Lines 013, 019, 019.5, 020, 021, 022.1 and 023If applicable, (make reference to the RMW revison) | The Product must have associated labelling that includes images and describes how the staining generated by the HER2 BMA CDx Assay can be interpreted, such that breast carcinoma specimens can be assigned a HER2 status (3+, 2+, 1+ and 0 with, and absent membrane staining).List all Product Requirements to be verified | HER2 BMA CDx Assay Interpretation Guide includes representative images of each HER2 status describes how staining of the HER2 BMA CDx Assay can be interpreted.Detail acceptance criteria for each of the requirements being verified or validated. Each acceptance criterion shall include sufficient detail to determine whether a requirement/need has been met without any ambiguity. | 13 | 12 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.07 | dRMW-016 Risk Lines 001, 007, 010 and 016 LBN-euRMW-020 Risk Lines 001, 003 009, 011 and 013 | The Product Instructions For Use (IFU) document must specify the use of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | 9 | 12 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.08 | LBN-dRMW-016 Risk Line 001, 007, 010 and 016 LBN euRMW-020 Risk Line 001 | The Product Instructions For Use (IFU) document must state any limitations relating to tissue fixation. | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | 5 | 12 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.18 | LBN-euRMW-020 Risk Lines 055 and 093 | The Product labelling will detail known limitations of formalin fixed, paraffin-embedded (FFPE) cut section stability. | Product Instructions For Use (IFU) document specifies the recommended storage conditions for cut tissue sections during the stability period (temperature, humidity, protection from light). | 2 | 12 | 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.22 | LBN-dRMW-016 Risk Line 034 LBN-euRMW-020 Risk Lines 078, 089 and 097 | The Product Instructions For Use (IFU) document must state a step-by-step outline of recommended procedures, including a validated immunohistochemistry (IHC) staining protocol and heat-induced epitope retrieval (HIER) protocol. | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | 4 | 12 | 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.38 | LBN-dRMW-016 Risk Lines 086, 088 and 089 LBN euRMW-020 Risk Line 095 | The Product Instructions For Use (IFU) document must advise on any limitations relating to use at altitude, temperature, or humidity. | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | 4 | 12 | 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.43 | LBN-dRMW-016 Risk Lines 071, 072, 073, 074 and 075 LBN-euRMW-020 Risk Line 044, 046 and 047 | The Product must have labeling that provides all required hazard and safety information and associated symbols. | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | 8 | 12 | 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.44 | LBN-dRMW-016 Risk Lines 078 and 081 | The symbols and text included on The Product labelling must be human legible. | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | 2 | 12 | 9 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.46 | LBN-dRMW-016 Risk Line 068 LBN-euRMW-020 Risk Lines 030, 033.5, 039, 044, 046,047, 091 and 095 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | 9 | 12 | 10 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.52 | N/A | The Product shall include compliance with UDI specifications. | The final labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | 0 | 12 | 11 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.53 | LBN-dRMW-016 Risk Line 071 | The Product labelling must be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | 1 | 12 | 12 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.54 | LBN-dRMW-016 Risk Line 071 | The Product labelling must be compliant with the relevant requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | 1 | 12 | 13 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.55 | LBN-dRMW-016 Risk Line 071 | The Product labelling must be compliant with the requirements of a Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | 1 | 12 | 14 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.56 | LBN-dRMW-016 Risk Line 071 | The Product labelling must be compliant with the relevant UK general IVD per the UK MDR 2002 labelling regulations. | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | 1 | 12 | 15 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.61 | LBN-dRMW-016 Risk Lines 010, 011, 016, 028, 039 and 040 LBN-euRMW-020 Risk Lines 019, 019.5, 020, 021, 022.1, 023 | The Product must include an associated scoring resource that provides examples of stained images and a pathologist proficiency test to guide competency expectations. | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | 12 | 12 | 16 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 Risk Lines 001, 007, 010, 016, 019, 020.5, 023, 024, 025, 028, 029, 030.1, 034, 042, 043, 046, 050, 053, 058, 059, 060, 062, 066, 068, 069, 084, 086, 087, 088 and 089 LBN-euRMW-020 Risks Lines 001, 003, 005, 009, 011, 011.5, 013, 015, 017, 044, 046, 047, 048, 049, 050, 052, 053, 055, 058, 064, 066, 068, 070, 072, 075, 078, 080, 083, 086, 089, 091, 092, 093, 095, 096, 097 and 098 | The Product Instructions for Use (IFU) and Interpretation Guide (IG) must contain all information specified as required IFU and IG content in the current Risk Control Measures defined in the Project Swift Risk Management Worksheets. | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | 67 | 12 | 17 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-003 | PR.63 | LBN-dRMW-016 Risk Lines 068, 071, 081 and 084 LBN-euRMW-020 Risks Lines 044 | The Product must have an associated Safety Data Sheet (SDS) and Hazard Notification that describe all required information. | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | 5 | 12 | 18 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52 | DP52-VEP-005 | PR.03 | dRMW-016 Risk Lines 001, 002, 003, 007, 010, 011 and 016 LBN-euRMW-020 Risk Lines 001, 003, 005, 009, 011 and 013TBC | The Product must be sensitive enough to detect HER2 protein expression in formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue across all HER2 expression levels (3+, 2+, 1+ and 0 with, and absent membrane staining). | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | 13 | 10 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-005 PREVALENCE Core Team Review.docx
DP52 | DP52-VEP-007 | PR.25 | dRMW-risk-029 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%.Point estimate of ≥90% when the lower bound of the 95% Confident Interval is ≥85%. | 0 | 16 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.26 | dRMW-risk-030.1 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance when different lots of detection kits are used | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | 0 | 16 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.27 | dRMW-risk-042 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance across different BOND III instruments (inter instrument) | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | 0 | 16 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.28 | dRMW-risk-043 (LBN-dRMW-016)2: The product does not maintain comparable and acceptable staining performance both inter staining runs on BOND-III (Inter Run) | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed within the same BOND Run Batch (inter-run repeatability). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | 0 | 16 | 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.29 | dRMW-risk-043 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance within BOND runs/SSAs (Intra run) | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining within BOND Runs (intra-run repeatability). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | 0 | 16 | 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-007 | PR.30 | dRMW-risk-046 (LBN-dRMW-016): The product does not maintain comparable and acceptable performance when used across different days (inter day) | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day repeatability). | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | 0 | 16 | 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52 | DP52-VEP-008 | PR.05 | dRMW-016 Risk Lines 003, 020.5, 025, 029, 030.1, 042, 043, 046, 050, 053, 058, 059, 060, 062, 066 and 069 LBN-euRMW-020 Risk Line 011.5, 024 and 072 | Product must be compatible with the HER2 BMA Cell Line Control. | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | 19 | 12 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-008 | PR.41 | euRMW-020 Risk Lines 093 | The Product must generate Acceptable Staining when there is a delay of up to seven days between placing The Product on the BOND III and commencement of staining. | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | 1 | 12 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-008 | PR.42 | euRMW-020 Risk Line 096 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining when there is a delay of up to 24 hours between completion of staining and removal. | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | 1 | 12 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-008 | PR.64 | LBN-dRMW-016 Risk Line 087 LBN-euRMW-020 Risks Line 053 | The Product must be compatible with reagent grade alcohol, xylene and xylene or toluene-based mounting media. | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | 2 | 12 | 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52 | DP52-VEP-010 | PR.06 | dRMW-016 Risk Lines 001, 002, 003, 007, 010, 011 and 016 LBN-euRMW-020 Risk Lines 001, 003, 005, 009, 011 and 013 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using required pre-analytic conditions. The Product, when used within the HER2 BMA Assay System, must generate Acceptable Staining in specimens of formalin fixed, paraffin-embedded (FFPE) breast. | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | 13 | 12 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-010 Sample Characterisation Ver Protocol.docx
DP52 | DP52-VEP-011 | PR.01 | dRMW-risk-001: Antibody is not specific/sensitive and does not detect HER2 in intended tissue dRMW-risk-002: Antibody is not specific and does not detect HER2 in intended tissue dRMW-risk-003: Antibody is not sensitive and does not detect HER2 in intended tissue dRMW-risk-007: Antibody is not specific/sensitive and does not detect HER2 in intended tissue dRMW-risk-010: Antibody is not specific and wrongly detects HER2 in intended tissue (background/non-specific staining present) dRMW-risk-011: Antibody is not specific and wrongly detects HER2 in intended tissue dRMW-risk-016: Antibody is not specific and wrongly detects HER2 in intended tissue. | The Product must be specific for the HER2 protein. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | 0 | 12 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52 | DP52-VEP-012 | PR.16 | LBN-dRMW-016 Risk Lines 058 and 059 LBN-euRMW-020 Risk Lines 042 and 043 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf-life (minimum 6-months shelf life with a minimum 2-months safety margin). | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 4 | 7 | 2 | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52 | DP52-VEP-012 | PR.10 | LBN-dRMW-016 Risk Line 060 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim | 1 | 7 | 3 | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52 | DP52-VEP-012 | PR.11 | LBN-dRMW-016 Risk Lines 042, 043, 046 and 069 LBN-euRMW-020 Risk Line 091 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine use conditions. | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | 5 | 7 | 4 | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52 | DP52-VEP-013 | PR.02 | LBN-dRMW-016 Risk Lines 023, 024, 025, 030.1 and 034 LBN-euRMW-020 Risk Lines 015, 017, 048, 049, 050, 052, 058, 064, 066 and 080dRMW-risk-025 | The Product must be compatible for use with all components of the HER2 BMA Assay System. | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | 16 | 11 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.12 | LBN-dRMW-016 Risk Line 023 and 024 LBN-euRMW-020 Risk Line 068 | The Product is only usable on the BOND-III instrument (not the BOND RX, BOND RXM, BOND-MAX, BOND-PRIME). | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | 3 | 11 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.13 | LBN-dRMW-016 Risk Lines 062 and 066 LBN-euRMW-020 Risk Line 092 | The Product must generate Acceptable Staining without having to be subjected to any preparation (e.g., warm-up, reconstitution, mixing, dilution, or titration), after removal from the recommended storage conditions (2-8°C). | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | 3 | 11 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.19 | LBN-dRMW-016 Risk Lines 024, 070 and 070.5dRMW-risk-070 dRMW-risk-070.5 dRMW-risk-080 | The Product must be correctly identified by the BOND-III fully automated system. | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | 4 | 11 | 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.20 | LBN-dRMW-016 Risk Lines 034 LBN-euRMW-020 Risk Lines 066 and 078dRMW-risk-034 | The immunohistochemistry (IHC) staining protocol for the HER2 BMA CDx Assay must be locked. | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | 4 | 11 | 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.21 | LBN-dRMW-016 Risk Line 086 LBN-euRMW-020 Risk Lines 001, 089 and 098dRMW-risk-086 dRMW-risk-034 dRMW-risk-088 dRMW-risk-089 dRMW-risk-093 | The heat-induced epitope retrieval (HIER) protocol for the HER2 BMA Assay must be unlocked. | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | 8 | 11 | 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.39 | LBN-dRMW-016 Risk Lines 056 and 057dRMW-risk-056 dRMW-risk-057 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance.The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. Stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | 2 | 11 | 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.40 | dRMW-risk-056 dRMW-risk-057 dRMW-risk-034N/A | The Product can be run within the same BOND Run Batch as other BOND Detection systems and staining protocol combinations. | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | 0 | 11 | 9 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.45 | LBN-dRMW-016 Risk Line 058 LBN-euRMW-020 Risk Line 042dRMW-risk-058 | The Product when expired, will not allow a staining run to commence. | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument.Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | 3 | 11 | 10 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 Risk Lines 001, 007, 010, 016, 020.5, 025, 029, 030.1, 042, 043, 046, 050, 053, 058, 059, 060, 062, 066 and 069 LBN-euRMW-020 Risk Lines 001, 003, 011.5, 021, 022.1, 042, 075, 080, 083, 089, 091, 095, 097 and 098 | The recommended tissue controls, when stained with The Product must demonstrate staining that will indicate BOND run validity. | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | 33 | 11 | 11 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP61 | DP61-VEP-002 | PR.06 | LBN-dRMW-017 Risk Line 013 LBN-euRMW-020 Risk Line 024 | The Product labelling must include images and describe how the product staining is interpreted. | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | 2 | 9 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.07 | LBN-dRMW-017 Risk Line 013 LBN-euRMW-020 Risk Line 024 | The Product labelling must describe the expected staining patterns and characteristics of each cell cores. | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | 2 | 9 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.23 | LBN-dRMW-017 Risk Line 028 | The Product must have labelling that provides all required hazard and safety information and associated symbols. | Users will review the Product slide mailer label, IFU and product specifications and complete a checklist to confirm that the required safety information is present. | 1 | 9 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.24 | LBN-dRMW-017 Risk Line 028 | The Product must not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices and the use of standard personal protective equipment (PPE). | Laboratory users will confirm that the hazards detailed in the IFU and product specifications can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP). | 1 | 9 | 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.25 | LBN-dRMW-017 Risk Line 032, 032.5, 033 and 033.5 LBN-euRMW-020 Risk Lines 035 and 037 LBN-dRMW-017 Risk lines 032, 032.5, 033, 033.5 050, 051 LBN-euRMW-020 Risk Line 035 and 037 | The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. The Immediate Packaging must be robust enough to ensure The Product remains intact and is safe to handle and use after routine shipment and storage. | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | 8 | 9 | 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.26 | LBN-dRMW-017 Risk Line 032, 032.5, 033 and 033.5 LBN-euRMW-020 Risk Lines and 035 | The outer packaging must be robust enough to ensure the containment of The Product if damaged or broken during routine shipping and storage. | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | 5 | 9 | 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.27 | LBN-euRMW-020 Risk Lines 027 and 035 | The Shipping Packaging must be robust enough to ensure The Product remains intact and is safe to handle after routine shipment and storage. | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | 2 | 9 | 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.28 | LBN-dRMW-017 Risk Lines 031 and 041 | The on-slide printed text must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired.The Product on-slide labels must be capable of withstanding shipping, storage within the recommended temperature range and normal use without their human legibility being impaired. | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. | 2 | 9 | 9 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.29 | LBN-dRMW-017 Risk Lines 043 and 044 | The Immediate Packaging label must be capable of withstanding shipping and storage within the recommended temperature range and normal use without becoming detached. | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remained fully adhered and undamaged. | 2 | 9 | 10 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.30 | LBN-dRMW-017 Risk Line 041 | The symbols and text included on The Product labelling must be human legible. | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. | 1 | 9 | 11 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.31 | N/A | The Product shall include compliance with UDI specifications. | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | 0 | 9 | 12 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.32 | LBN-dRMW-017 Risk Line 029 | The Product labelling must be created to be compliant with US regulation 21 CFR Part 801 and 809 - Labeling for in vitro diagnostic products. | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. | 1 | 9 | 13 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.33 | LBN-dRMW-017 Risk Line 029 | The Product labelling must be created to be compliant with the relevant requirements from IVDR EU/2017/746. | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | 1 | 9 | 14 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.34 | LBN-dRMW-017 Risk Line 029 | The Product labelling must be created to be compliant with the relevant Canadian IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | 1 | 9 | 15 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.35 | LBN-dRMW-017 Risk Line 029 | The Product labelling must be created to comply with the relevant UK general IVD as per the UK MDR 2002 labelling regulations. | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. | 1 | 9 | 16 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 Risk Lines 006, 011, 013, 028, 033, 033.5 and 049 LBN-euRMW-020 Risk Line 072, 075, 091 and 095 | The Product Instruction For Use (IFU) and Interpretation Guide (IG) must contain all required information specified as IFU and IG content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets. | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | 11 | 9 | 17 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61 | DP61-VEP-003 | PR.13 | LBN-dRMW-017 Risk Lines 021 and 039 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of The Product are varied (inter-lot repeatability). | Stained CLCs must demonstrate Acceptable Staining and Consistent Staining across six different lots. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.For inter-lot testing, the lower bound of the two-sided 95% confidence interval for the overall percent agreement should be at least 85%. Testing will be performed on 6 lots. | 2 | 10 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.14 | LBN-dRMW-017 Risk Line 022 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining throughout the same lots of The Product (intra-lot repeatability). | Stained CLCs For intra-lot testing, taken from within a single lot must generate Acceptable Staining and Consistent Staining. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate the overall percent agreement should be at least 85%. Only sections within a single lot will be tested. | 1 | 10 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.15 | LBN-dRMW-017 Risk Line 022 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different non-consecutive days (inter-day reproducibility). | When stained over 5 non-consecutive days (spanning a minimum period of 20 days), CLCs must generate Acceptable Staining and Consistent Staining. For inter-day testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate for the overall percent agreement should should be at least 85%. Testing will occur over 5 non-consecutive days spanning a minimum period of 20 days. | 1 | 10 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.16 | LBN-dRMW-017 Risk Line 022 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across different BOND-III fully automated systems (inter-instrument repeatability). | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across three different BOND-IIIs. For inter-instrument testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed across three BOND-IIIs. | 1 | 10 | 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.17 | N/A | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining across BOND Runs performed both within BOND Runs andwithin the across BOND Runs performed within the same BOND Run Batch (interintra-run and inter-run repeatability). | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across different SSAs on a BOND-III Instrument run batch and within BOND Runs (SSAs). For inter-run testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Iner-run data will be collected across BOND run batches. | 0 | 10 | 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.18 | LBN-dRMW-017 Risk Line 021.2 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Polymer Refine Detection (DS9800) are varied (inter-detection lot repeatability). | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system For inter-detectiontested samples lot repeatability testing, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Three lots will be used during testing. | 1 | 10 | 7 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-003 | PR.19 | LBN-dRMW-017 Risk Line 021.1 | The Product, when used within the HER2 BMA CDx Assay System, must generate Acceptable Staining and Consistent Staining when lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx antibody are varied (inter-antibody lot repeatability). | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. For inter-antibody lot repeatability testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed using 3 lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. | 1 | 10 | 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-005 | PR.08 | LBN-dRMW-017 Risk Line 018 | The Product must be capable of performing as intended and generating Acceptable Staining after being subjected to routine transport conditions. | 100% of stained Product must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide after being subject to transport simulation.The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | 1 | 12 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61 | DP61-VEP-005 | PR.11 | LBN-dRMW-017 Risk Line 017 LBN-euRMW-020 Risk Line 043 | The Product must be capable of performing as intended and generating Acceptable Staining for the duration of its shelf life when stored in the immediate packaging (minimum of 18 months shelf life with a minimum 3 month safety margin). | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.Following a minimum of 18 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The Product must also retain its ability to correctly detect procedural failures. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | 2 | 12 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61 | DP61-VEP-005 | PR.12 | LBN-dRMW-017 Risk Line 039 | The Intermediate Product must be stable for a minimum of six months when stored at room temperature. | Following a minimum of six months storage at room temperature, 100% of The Intermediate Product (individual agarose cell line cores of BT474-M3, RT4 and PANC-1) continues to produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1 must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide after being stored at room temperature for at least six months. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | 1 | 12 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61 | DP61-VEP-006 | PR.04 | LBN-dRMW-017 Risk Line 009 | The Product must be compatible for use with all components of the HER2 BMA CDx Assay System. | When stained using all components of the HER2 BMA CDx Assay system, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using all components of the HER2 BMA CDx Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | 1 | 10 | 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.05 | LBN-dRMW-017 Risk Line 009 | The Product must be compatible for use on the BOND-III fully automated system. | The Product, when loaded onto the BOND-III instrument, must not trigger any warning or error messages upon initiation of a staining run. The Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | 1 | 10 | 3 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.09 | LBN-dRMW-017 Risk Lines 001 and , 004, 019 and 026 | The Product must generate Acceptable Staining without having to be subjected to any immediate preparation (e.g., slide baking or incubation) after being removed from the recommended storage conditions (2-8°C). | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | 4 | 10 | 4 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.10 | LBN-dRMW-017 Risk Line 026N/A | The Product must generate Acceptable Staining when used in conjunction with the recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols. | When stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | 1 | 10 | 5 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61 | DP61-VEP-006 | PR.22 | LBN-dRMW-017 Risk Line 006 | The placement of the cell line cores within The Product must be consistently orientated relative to each other. | A HER2 scoring SME/qualified pathologist will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product IFU and HER2 BMA CDx Assay Interpretation Guide. 100% of assessed CLC slides must demonstrate correct core orientation. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide. | 1 | 10 | 6 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP52 | DP52/DP61-VEP-001 (Environmental) | PR.35 | LBN-dRMW-016 Risk Line 086; LBN-euRMW-020 Risk Line 098 | The Product, when used within the BOND-III altitude specifications. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | 2 | Table 15 | Page 27 | C:\Users\emma.goodrick\Downloads\DP5261-VEP-001_-_Environmental_Verification_Protocol_for_the_HER2_BMA_CDx_Assay Signed 28-July-2026.pdf
DP61 | DP52/DP61-VEP-001 (Environmental) | PR.36 | LBN-dRMW-017 Risk Line 049 | The Product, when used within the BOND-III altitude specifications. | To pass, when stained within the BOND-III altitude specifications (0-1600m) using either the default HIER or alternative HIERs (ER1, 97°C for 35 mins, and ER1, 96°C for 25 mins), 100% of sections stained should demonstrate no evidence of altitude staining artifact impacting slide interpretation as defined in Section 5.3.1. Additionally, the HER2 status for HER2 slides (CLCs/HER2 test tissue and positive/negative control samples) stained at each simulated altitude when compared to the sea-level control must meet the lower limit of the two-sided 95% confidence interval for the pass rate >=85%. | 1 | Table 15 | Page 27 | C:\Users\emma.goodrick\Downloads\DP5261-VEP-001_-_Environmental_Verification_Protocol_for_the_HER2_BMA_CDx_Assay Signed 28-July-2026.pdf
DP52 | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | PR.14 | N/A - RP101 uses generic BRTU risk-group identifiers; project-specific links are in DP52-PR-001 | An individual 7 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 46 slides when operated in 150 μl dispense mode. | The product has sufficient volume to stain 46 slides (7 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 46 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | 0 | RP101-IUVP-001 Section 10; RP101-IUVR-001 Section 5 | Page Plan page 16; report page 11 | T:\R&D\PROJECTS\NPD Overarching Documents\In-Use Verification\RP101-IUVP-001. NPD Overarching In-Use Verification plan.pdf
DP52 | RP101-IUVP-001 / RP101-IUVR-001 (In-Use) | PR.15 | N/A - RP101 uses generic BRTU risk-group identifiers; project-specific links are in DP52-PR-001 | An individual 30 mL unit of the Product must include sufficient reagent volume to allow staining of a minimum of 200 slides when operated in 150 μl dispense mode. | The product has sufficient volume to stain 200 slides (30 mL format) when the BOND-III is operated in single 150 μL dispense mode. The product will FAIL if the BOND-III is unable to dispense onto all 200 slides. As this study is being used to mitigate risks for future new product development (NPD), a success rate of 100% is required for each product on each platform. For the overall study to pass, 100% of the products tested must meet the above criteria. | 0 | RP101-IUVP-001 Section 10; RP101-IUVR-001 Section 5 | Page Plan page 16; report page 11 | T:\R&D\PROJECTS\NPD Overarching Documents\In-Use Verification\RP101-IUVP-001. NPD Overarching In-Use Verification plan.pdf

## Protocol Risk Refs (hidden)
ProductFamily | ProtocolID | ProductRequirementID | RiskSourceDocument | RiskID | RiskExistsInWorkbook | RiskWorkbookExcelRow | VerificationStatus | SourceTable | SourceTableRow
DP52 | DP52-VEP-003 | PR.04 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-dRMW-016 | dRMW-risk-028 | Yes | 23 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-dRMW-016 | dRMW-risk-039 | Yes | 27 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-dRMW-016 | dRMW-risk-040 | Yes | 28 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-euRMW-020 | euRMW-risk-019 | Yes | 18 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-euRMW-020 | euRMW-risk-019.5 | Yes | 19 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-euRMW-020 | euRMW-risk-020 | Yes | 20 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-euRMW-020 | euRMW-risk-021 | Yes | 21 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-euRMW-020 | euRMW-risk-022.1 | Yes | 22 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.04 | LBN-euRMW-020 | euRMW-risk-023 | Yes | 23 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-003 | PR.07 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-euRMW-020 | euRMW-risk-001 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-euRMW-020 | euRMW-risk-009 | Yes | 10 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-euRMW-020 | euRMW-risk-011 | Yes | 11 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.07 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-003 | PR.08 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | Verified exact ID exists | 12 | 4
DP52 | DP52-VEP-003 | PR.08 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | Verified exact ID exists | 12 | 4
DP52 | DP52-VEP-003 | PR.08 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 12 | 4
DP52 | DP52-VEP-003 | PR.08 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 12 | 4
DP52 | DP52-VEP-003 | PR.08 | LBN-euRMW-020 | euRMW-risk-001 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 4
DP52 | DP52-VEP-003 | PR.18 | LBN-euRMW-020 | euRMW-risk-055 | Yes | 40 | Verified exact ID exists | 12 | 5
DP52 | DP52-VEP-003 | PR.18 | LBN-euRMW-020 | euRMW-risk-093 | Yes | 55 | Verified exact ID exists | 12 | 5
DP52 | DP52-VEP-003 | PR.22 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | Verified exact ID exists | 12 | 6
DP52 | DP52-VEP-003 | PR.22 | LBN-euRMW-020 | euRMW-risk-078 | Yes | 48 | Verified exact ID exists | 12 | 6
DP52 | DP52-VEP-003 | PR.22 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | Verified exact ID exists | 12 | 6
DP52 | DP52-VEP-003 | PR.22 | LBN-euRMW-020 | euRMW-risk-097 | Yes | 57 | Verified exact ID exists | 12 | 6
DP52 | DP52-VEP-003 | PR.38 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | Verified exact ID exists | 12 | 7
DP52 | DP52-VEP-003 | PR.38 | LBN-dRMW-016 | dRMW-risk-088 | Yes | 59 | Verified exact ID exists | 12 | 7
DP52 | DP52-VEP-003 | PR.38 | LBN-dRMW-016 | dRMW-risk-089 | Yes | 60 | Verified exact ID exists | 12 | 7
DP52 | DP52-VEP-003 | PR.38 | LBN-euRMW-020 | euRMW-risk-095 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 7
DP52 | DP52-VEP-003 | PR.43 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.43 | LBN-dRMW-016 | dRMW-risk-072 | Yes | 46 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.43 | LBN-dRMW-016 | dRMW-risk-073 | Yes | 47 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.43 | LBN-dRMW-016 | dRMW-risk-074 | Yes | 48 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.43 | LBN-dRMW-016 | dRMW-risk-075 | Yes | 49 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.43 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.43 | LBN-euRMW-020 | euRMW-risk-046 | Yes | 33 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.43 | LBN-euRMW-020 | euRMW-risk-047 | Yes | 34 | Verified exact ID exists | 12 | 8
DP52 | DP52-VEP-003 | PR.44 | LBN-dRMW-016 | dRMW-risk-078 | Yes | 51 | Verified exact ID exists | 12 | 9
DP52 | DP52-VEP-003 | PR.44 | LBN-dRMW-016 | dRMW-risk-081 | Yes | 53 | Verified exact ID exists | 12 | 9
DP52 | DP52-VEP-003 | PR.46 | LBN-dRMW-016 | dRMW-risk-068 | Yes | 41 | Verified exact ID exists | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-030 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-033.5 | Yes | 27 | Verified exact ID exists | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-039 | Yes | 29 | Verified exact ID exists | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | Verified exact ID exists | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-046 | Yes | 33 | Verified exact ID exists | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-047 | Yes | 34 | Verified exact ID exists | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | Verified exact ID exists | 12 | 10
DP52 | DP52-VEP-003 | PR.46 | LBN-euRMW-020 | euRMW-risk-095 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 10
DP52 | DP52-VEP-003 | PR.53 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | Verified exact ID exists | 12 | 12
DP52 | DP52-VEP-003 | PR.54 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | Verified exact ID exists | 12 | 13
DP52 | DP52-VEP-003 | PR.55 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | Verified exact ID exists | 12 | 14
DP52 | DP52-VEP-003 | PR.56 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | Verified exact ID exists | 12 | 15
DP52 | DP52-VEP-003 | PR.61 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-dRMW-016 | dRMW-risk-028 | Yes | 23 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-dRMW-016 | dRMW-risk-039 | Yes | 27 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-dRMW-016 | dRMW-risk-040 | Yes | 28 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-euRMW-020 | euRMW-risk-019 | Yes | 18 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-euRMW-020 | euRMW-risk-019.5 | Yes | 19 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-euRMW-020 | euRMW-risk-020 | Yes | 20 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-euRMW-020 | euRMW-risk-021 | Yes | 21 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-euRMW-020 | euRMW-risk-022.1 | Yes | 22 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.61 | LBN-euRMW-020 | euRMW-risk-023 | Yes | 23 | Verified exact ID exists | 12 | 16
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-019 | Yes | 15 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-023 | Yes | 20 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-028 | Yes | 23 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-029 | Yes | 24 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-050 | Yes | 32 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-053 | Yes | 33 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-068 | Yes | 41 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-084 | Yes | 55 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-087 | Yes | 58 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-088 | Yes | 59 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-dRMW-016 | dRMW-risk-089 | Yes | 60 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-001 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-005 | Yes | 9 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-009 | Yes | 10 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-011 | Yes | 11 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-011.5 | Yes | 14 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-015 | Yes | 16 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-017 | Yes | 17 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-046 | Yes | 33 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-047 | Yes | 34 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-048 | Yes | 35 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-049 | Yes | 36 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-050 | Yes | 37 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-052 | Yes | 38 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-053 | Yes | 39 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-055 | Yes | 40 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-058 | Yes | 41 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-064 | Yes | 42 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-066 | Yes | 43 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-068 | Yes | 44 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-070 | Yes | 45 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-072 | Yes | 46 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-075 | Yes | 47 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-078 | Yes | 48 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-080 | Yes | 49 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-083 | Yes | 50 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-086 | Yes | 51 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-092 | Yes | 54 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-093 | Yes | 55 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-095 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-096 | Yes | 56 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-097 | Yes | 57 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.62 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | Verified exact ID exists | 12 | 17
DP52 | DP52-VEP-003 | PR.63 | LBN-dRMW-016 | dRMW-risk-068 | Yes | 41 | Verified exact ID exists | 12 | 18
DP52 | DP52-VEP-003 | PR.63 | LBN-dRMW-016 | dRMW-risk-071 | Yes | 45 | Verified exact ID exists | 12 | 18
DP52 | DP52-VEP-003 | PR.63 | LBN-dRMW-016 | dRMW-risk-081 | Yes | 53 | Verified exact ID exists | 12 | 18
DP52 | DP52-VEP-003 | PR.63 | LBN-dRMW-016 | dRMW-risk-084 | Yes | 55 | Verified exact ID exists | 12 | 18
DP52 | DP52-VEP-003 | PR.63 | LBN-euRMW-020 | euRMW-risk-044 | Yes | 32 | Verified exact ID exists | 12 | 18
DP52 | DP52-VEP-005 | PR.03 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-dRMW-016 | dRMW-risk-002 | Yes | 9 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-dRMW-016 | dRMW-risk-003 | Yes | 10 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-euRMW-020 | euRMW-risk-001 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-euRMW-020 | euRMW-risk-005 | Yes | 9 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-euRMW-020 | euRMW-risk-009 | Yes | 10 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-euRMW-020 | euRMW-risk-011 | Yes | 11 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-005 | PR.03 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | Verified exact ID exists | 10 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-003 | Yes | 10 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-029 | Yes | 24 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-050 | Yes | 32 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-053 | Yes | 33 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-euRMW-020 | euRMW-risk-011.5 | Yes | 14 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.05 | LBN-euRMW-020 | euRMW-risk-072 | Yes | 46 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-008 | PR.41 | LBN-euRMW-020 | euRMW-risk-093 | Yes | 55 | Verified exact ID exists | 12 | 3
DP52 | DP52-VEP-008 | PR.42 | LBN-euRMW-020 | euRMW-risk-096 | Yes | 56 | Verified exact ID exists | 12 | 4
DP52 | DP52-VEP-008 | PR.64 | LBN-dRMW-016 | dRMW-risk-087 | Yes | 58 | Verified exact ID exists | 12 | 5
DP52 | DP52-VEP-008 | PR.64 | LBN-euRMW-020 | euRMW-risk-053 | Yes | 39 | Verified exact ID exists | 12 | 5
DP52 | DP52-VEP-010 | PR.06 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-dRMW-016 | dRMW-risk-002 | Yes | 9 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-dRMW-016 | dRMW-risk-003 | Yes | 10 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-dRMW-016 | dRMW-risk-011 | Yes | 13 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-euRMW-020 | euRMW-risk-001 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-euRMW-020 | euRMW-risk-005 | Yes | 9 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-euRMW-020 | euRMW-risk-009 | Yes | 10 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-euRMW-020 | euRMW-risk-011 | Yes | 11 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-010 | PR.06 | LBN-euRMW-020 | euRMW-risk-013 | Yes | 15 | Verified exact ID exists | 12 | 2
DP52 | DP52-VEP-012 | PR.16 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | Verified exact ID exists | 7 | 2
DP52 | DP52-VEP-012 | PR.16 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | Verified exact ID exists | 7 | 2
DP52 | DP52-VEP-012 | PR.16 | LBN-euRMW-020 | euRMW-risk-042 | Yes | 30 | Verified exact ID exists | 7 | 2
DP52 | DP52-VEP-012 | PR.16 | LBN-euRMW-020 | euRMW-risk-043 | Yes | 31 | Verified exact ID exists | 7 | 2
DP52 | DP52-VEP-012 | PR.10 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | Verified exact ID exists | 7 | 3
DP52 | DP52-VEP-012 | PR.11 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | Verified exact ID exists | 7 | 4
DP52 | DP52-VEP-012 | PR.11 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | Verified exact ID exists | 7 | 4
DP52 | DP52-VEP-012 | PR.11 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | Verified exact ID exists | 7 | 4
DP52 | DP52-VEP-012 | PR.11 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | Verified exact ID exists | 7 | 4
DP52 | DP52-VEP-012 | PR.11 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | Verified exact ID exists | 7 | 4
DP52 | DP52-VEP-013 | PR.02 | LBN-dRMW-016 | dRMW-risk-023 | Yes | 20 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-015 | Yes | 16 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-017 | Yes | 17 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-025 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-048 | Yes | 35 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-049 | Yes | 36 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-050 | Yes | 37 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-052 | Yes | 38 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-058 | Yes | 41 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-064 | Yes | 42 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-066 | Yes | 43 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.02 | LBN-euRMW-020 | euRMW-risk-080 | Yes | 49 | Verified exact ID exists | 11 | 2
DP52 | DP52-VEP-013 | PR.12 | LBN-dRMW-016 | dRMW-risk-023 | Yes | 20 | Verified exact ID exists | 11 | 3
DP52 | DP52-VEP-013 | PR.12 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | Verified exact ID exists | 11 | 3
DP52 | DP52-VEP-013 | PR.12 | LBN-euRMW-020 | euRMW-risk-068 | Yes | 44 | Verified exact ID exists | 11 | 3
DP52 | DP52-VEP-013 | PR.13 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | Verified exact ID exists | 11 | 4
DP52 | DP52-VEP-013 | PR.13 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | Verified exact ID exists | 11 | 4
DP52 | DP52-VEP-013 | PR.13 | LBN-euRMW-020 | euRMW-risk-092 | Yes | 54 | Verified exact ID exists | 11 | 4
DP52 | DP52-VEP-013 | PR.19 | LBN-dRMW-016 | dRMW-risk-024 | Yes | 21 | Verified exact ID exists | 11 | 5
DP52 | DP52-VEP-013 | PR.19 | LBN-dRMW-016 | dRMW-risk-070 | Yes | 43 | Verified exact ID exists | 11 | 5
DP52 | DP52-VEP-013 | PR.19 | LBN-dRMW-016 | dRMW-risk-070.5 | Yes | 44 | Verified exact ID exists | 11 | 5
DP52 | DP52-VEP-013 | PR.19 | LBN-dRMW-016 | dRMW-risk-080 | Yes | 52 | Verified exact ID exists | 11 | 5
DP52 | DP52-VEP-013 | PR.20 | LBN-dRMW-016 | dRMW-risk-034 | Yes | 26 | Verified exact ID exists | 11 | 6
DP52 | DP52-VEP-013 | PR.20 | LBN-euRMW-020 | euRMW-risk-034 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 6
DP52 | DP52-VEP-013 | PR.20 | LBN-euRMW-020 | euRMW-risk-066 | Yes | 43 | Verified exact ID exists | 11 | 6
DP52 | DP52-VEP-013 | PR.20 | LBN-euRMW-020 | euRMW-risk-078 | Yes | 48 | Verified exact ID exists | 11 | 6
DP52 | DP52-VEP-013 | PR.21 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | Verified exact ID exists | 11 | 7
DP52 | DP52-VEP-013 | PR.21 | LBN-euRMW-020 | euRMW-risk-001 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 7
DP52 | DP52-VEP-013 | PR.21 | LBN-euRMW-020 | euRMW-risk-034 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 7
DP52 | DP52-VEP-013 | PR.21 | LBN-euRMW-020 | euRMW-risk-086 | Yes | 51 | Verified exact ID exists | 11 | 7
DP52 | DP52-VEP-013 | PR.21 | LBN-euRMW-020 | euRMW-risk-088 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 7
DP52 | DP52-VEP-013 | PR.21 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | Verified exact ID exists | 11 | 7
DP52 | DP52-VEP-013 | PR.21 | LBN-euRMW-020 | euRMW-risk-093 | Yes | 55 | Verified exact ID exists | 11 | 7
DP52 | DP52-VEP-013 | PR.21 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | Verified exact ID exists | 11 | 7
DP52 | DP52-VEP-013 | PR.39 | LBN-dRMW-016 | dRMW-risk-056 | Yes | 34 | Verified exact ID exists | 11 | 8
DP52 | DP52-VEP-013 | PR.39 | LBN-dRMW-016 | dRMW-risk-057 | Yes | 35 | Verified exact ID exists | 11 | 8
DP52 | DP52-VEP-013 | PR.45 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | Verified exact ID exists | 11 | 10
DP52 | DP52-VEP-013 | PR.45 | LBN-euRMW-020 | euRMW-risk-042 | Yes | 30 | Verified exact ID exists | 11 | 10
DP52 | DP52-VEP-013 | PR.45 | LBN-euRMW-020 | euRMW-risk-058 | Yes | 41 | Verified exact ID exists | 11 | 10
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-001 | Yes | 8 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-007 | Yes | 11 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-010 | Yes | 12 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-016 | Yes | 14 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-020.5 | Yes | 17 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-025 | Yes | 22 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-029 | Yes | 24 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-030.1 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-042 | Yes | 29 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-043 | Yes | 30 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-046 | Yes | 31 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-050 | Yes | 32 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-053 | Yes | 33 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-058 | Yes | 36 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-059 | Yes | 37 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-060 | Yes | 38 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-062 | Yes | 39 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-066 | Yes | 40 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-dRMW-016 | dRMW-risk-069 | Yes | 42 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-001 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-003 | Yes | 8 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-011.5 | Yes | 14 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-021 | Yes | 21 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-022.1 | Yes | 22 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-042 | Yes | 30 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-075 | Yes | 47 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-080 | Yes | 49 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-083 | Yes | 50 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-089 | Yes | 52 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-095 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-097 | Yes | 57 | Verified exact ID exists | 11 | 11
DP52 | DP52-VEP-013 | PR.59 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | Verified exact ID exists | 11 | 11
DP61 | DP61-VEP-002 | PR.06 | LBN-dRMW-017 | dRMW-risk-013 | Yes | 15 | Verified exact ID exists | 9 | 2
DP61 | DP61-VEP-002 | PR.06 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | Verified exact ID exists | 9 | 2
DP61 | DP61-VEP-002 | PR.07 | LBN-dRMW-017 | dRMW-risk-013 | Yes | 15 | Verified exact ID exists | 9 | 3
DP61 | DP61-VEP-002 | PR.07 | LBN-euRMW-020 | euRMW-risk-024 | Yes | 24 | Verified exact ID exists | 9 | 3
DP61 | DP61-VEP-002 | PR.23 | LBN-dRMW-017 | dRMW-risk-028 | Yes | 22 | Verified exact ID exists | 9 | 4
DP61 | DP61-VEP-002 | PR.24 | LBN-dRMW-017 | dRMW-risk-028 | Yes | 22 | Verified exact ID exists | 9 | 5
DP61 | DP61-VEP-002 | PR.25 | LBN-dRMW-017 | dRMW-risk-032 | Yes | 26 | Verified exact ID exists | 9 | 6
DP61 | DP61-VEP-002 | PR.25 | LBN-dRMW-017 | dRMW-risk-032.5 | Yes | 27 | Verified exact ID exists | 9 | 6
DP61 | DP61-VEP-002 | PR.25 | LBN-dRMW-017 | dRMW-risk-033 | Yes | 28 | Verified exact ID exists | 9 | 6
DP61 | DP61-VEP-002 | PR.25 | LBN-dRMW-017 | dRMW-risk-033.5 | Yes | 29 | Verified exact ID exists | 9 | 6
DP61 | DP61-VEP-002 | PR.25 | LBN-dRMW-017 | dRMW-risk-050 | Yes | 44 | Verified exact ID exists | 9 | 6
DP61 | DP61-VEP-002 | PR.25 | LBN-dRMW-017 | dRMW-risk-051 | Yes | 45 | Verified exact ID exists | 9 | 6
DP61 | DP61-VEP-002 | PR.25 | LBN-euRMW-020 | euRMW-risk-035 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 9 | 6
DP61 | DP61-VEP-002 | PR.25 | LBN-euRMW-020 | euRMW-risk-037 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 9 | 6
DP61 | DP61-VEP-002 | PR.26 | LBN-dRMW-017 | dRMW-risk-032 | Yes | 26 | Verified exact ID exists | 9 | 7
DP61 | DP61-VEP-002 | PR.26 | LBN-dRMW-017 | dRMW-risk-032.5 | Yes | 27 | Verified exact ID exists | 9 | 7
DP61 | DP61-VEP-002 | PR.26 | LBN-dRMW-017 | dRMW-risk-033 | Yes | 28 | Verified exact ID exists | 9 | 7
DP61 | DP61-VEP-002 | PR.26 | LBN-dRMW-017 | dRMW-risk-033.5 | Yes | 29 | Verified exact ID exists | 9 | 7
DP61 | DP61-VEP-002 | PR.26 | LBN-euRMW-020 | euRMW-risk-035 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 9 | 7
DP61 | DP61-VEP-002 | PR.27 | LBN-euRMW-020 | euRMW-risk-027 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 9 | 8
DP61 | DP61-VEP-002 | PR.27 | LBN-euRMW-020 | euRMW-risk-035 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 9 | 8
DP61 | DP61-VEP-002 | PR.28 | LBN-dRMW-017 | dRMW-risk-031 | Yes | 25 | Verified exact ID exists | 9 | 9
DP61 | DP61-VEP-002 | PR.28 | LBN-dRMW-017 | dRMW-risk-041 | Yes | 33 | Verified exact ID exists | 9 | 9
DP61 | DP61-VEP-002 | PR.29 | LBN-dRMW-017 | dRMW-risk-043 | Yes | 35 | Verified exact ID exists | 9 | 10
DP61 | DP61-VEP-002 | PR.29 | LBN-dRMW-017 | dRMW-risk-044 | Yes | 36 | Verified exact ID exists | 9 | 10
DP61 | DP61-VEP-002 | PR.30 | LBN-dRMW-017 | dRMW-risk-041 | Yes | 33 | Verified exact ID exists | 9 | 11
DP61 | DP61-VEP-002 | PR.32 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | Verified exact ID exists | 9 | 13
DP61 | DP61-VEP-002 | PR.33 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | Verified exact ID exists | 9 | 14
DP61 | DP61-VEP-002 | PR.34 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | Verified exact ID exists | 9 | 15
DP61 | DP61-VEP-002 | PR.35 | LBN-dRMW-017 | dRMW-risk-029 | Yes | 23 | Verified exact ID exists | 9 | 16
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 | dRMW-risk-006 | Yes | 10 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 | dRMW-risk-011 | Yes | 14 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 | dRMW-risk-013 | Yes | 15 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 | dRMW-risk-028 | Yes | 22 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 | dRMW-risk-033 | Yes | 28 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 | dRMW-risk-033.5 | Yes | 29 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-dRMW-017 | dRMW-risk-049 | Yes | 41 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-euRMW-020 | euRMW-risk-072 | Yes | 46 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-euRMW-020 | euRMW-risk-075 | Yes | 47 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-euRMW-020 | euRMW-risk-091 | Yes | 53 | Verified exact ID exists | 9 | 17
DP61 | DP61-VEP-002 | PR.37 | LBN-euRMW-020 | euRMW-risk-095 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 9 | 17
DP61 | DP61-VEP-003 | PR.13 | LBN-dRMW-017 | dRMW-risk-021 | Yes | 18 | Verified exact ID exists | 10 | 2
DP61 | DP61-VEP-003 | PR.13 | LBN-dRMW-017 | dRMW-risk-039 | Yes | 31 | Verified exact ID exists | 10 | 2
DP61 | DP61-VEP-003 | PR.14 | LBN-dRMW-017 | dRMW-risk-022 | Yes | 21 | Verified exact ID exists | 10 | 3
DP61 | DP61-VEP-003 | PR.15 | LBN-dRMW-017 | dRMW-risk-022 | Yes | 21 | Verified exact ID exists | 10 | 4
DP61 | DP61-VEP-003 | PR.16 | LBN-dRMW-017 | dRMW-risk-022 | Yes | 21 | Verified exact ID exists | 10 | 5
DP61 | DP61-VEP-003 | PR.18 | LBN-dRMW-017 | dRMW-risk-021.2 | Yes | 20 | Verified exact ID exists | 10 | 7
DP61 | DP61-VEP-003 | PR.19 | LBN-dRMW-017 | dRMW-risk-021.1 | Yes | 19 | Verified exact ID exists | 10 | 8
DP61 | DP61-VEP-005 | PR.08 | LBN-dRMW-017 | dRMW-risk-018 | Yes | 17 | Verified exact ID exists | 12 | 2
DP61 | DP61-VEP-005 | PR.11 | LBN-dRMW-017 | dRMW-risk-017 | Yes | 16 | Verified exact ID exists | 12 | 3
DP61 | DP61-VEP-005 | PR.11 | LBN-euRMW-020 | euRMW-risk-043 | Yes | 31 | Verified exact ID exists | 12 | 3
DP61 | DP61-VEP-005 | PR.12 | LBN-dRMW-017 | dRMW-risk-039 | Yes | 31 | Verified exact ID exists | 12 | 4
DP61 | DP61-VEP-006 | PR.04 | LBN-dRMW-017 | dRMW-risk-009 | Yes | 12 | Verified exact ID exists | 10 | 2
DP61 | DP61-VEP-006 | PR.05 | LBN-dRMW-017 | dRMW-risk-009 | Yes | 12 | Verified exact ID exists | 10 | 3
DP61 | DP61-VEP-006 | PR.09 | LBN-dRMW-017 | dRMW-risk-001 | Yes | 8 | Verified exact ID exists | 10 | 4
DP61 | DP61-VEP-006 | PR.09 | LBN-dRMW-017 | dRMW-risk-004 | Yes | 9 | Verified exact ID exists | 10 | 4
DP61 | DP61-VEP-006 | PR.09 | LBN-dRMW-017 | dRMW-risk-019 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 10 | 4
DP61 | DP61-VEP-006 | PR.09 | LBN-dRMW-017 | dRMW-risk-026 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 10 | 4
DP61 | DP61-VEP-006 | PR.10 | LBN-dRMW-017 | dRMW-risk-026 | No |  | [NEEDS INPUT] Referenced ID not found in supplied risk workbook | 10 | 5
DP61 | DP61-VEP-006 | PR.22 | LBN-dRMW-017 | dRMW-risk-006 | Yes | 10 | Verified exact ID exists | 10 | 6
DP52 | DP52/DP61-VEP-001 (Environmental) | PR.35 | LBN-dRMW-016 | dRMW-risk-086 | Yes | 57 | Verified exact ID exists | Table 15 | Page 27
DP52 | DP52/DP61-VEP-001 (Environmental) | PR.35 | LBN-euRMW-020 | euRMW-risk-098 | Yes | 58 | Verified exact ID exists | Table 15 | Page 27
DP61 | DP52/DP61-VEP-001 (Environmental) | PR.36 | LBN-dRMW-017 | dRMW-risk-049 | Yes | 41 | Verified exact ID exists | Table 15 | Page 27
DP61 | DP52/DP61-VEP-001 (Environmental) | N/A | LBN-dRMW-017 | dRMW-risk-049.01 | Yes | 42 | Verified exact ID exists | Table 15 | Page 27-28
DP61 | DP52/DP61-VEP-001 (Environmental) | N/A | LBN-dRMW-017 | dRMW-risk-049.1 | Yes | 43 | Verified exact ID exists | Table 15 | Page 28

## Study Criteria (hidden)
ProtocolID | SourceTable | SourceTableRow | StudyOrParameter | AcceptanceCriteria | SourceFile
DP61-VEP-001 | 5 | 2 | PR.01 | The Product must be able to confirm that assay reagents, instruments, and protocols are operating correctly during staining runs. Deviations in the assay reagents, instrument and protocol parameters should result in a failure to meet the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 3 | PR.02 | The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA CDx, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression.The Product must include cell line cores that, when stained with the BOND Ready-to-Use Primary Antibody HER2 BMA, demonstrate no HER2 expression, low HER2 expression, and high HER2 expression. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 4 | PR.03 | Deviations in the assay reagents, instrument and protocol that alter the HER2 expression status in patient samples should result in a failure of the CLC to meet the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 5 | PR.04 | The Product, when stained using all components of the HER2 BMA Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 6 | PR.05 | The Product, when loaded onto the BOND-III instrument, must not trigger any warning or error messages upon initiation of a staining run. The Product must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 7 | PR.06 | HER2 BMA CDx Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to ensure it includes representative images and clearly describes how staining of the Cell Line Control can be interpreted. Sample Size: TBD1 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 8 | PR.07 | HER2 BMA Assay Interpretation Guide and Instructions For Use (IFU) to be checked by HER2 scoring SMEs/qualified pathologists to confirm they describes the expected staining patterns and characteristics of each cell line core. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 9 | PR.08 | The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 10 | PR.09 | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 11 | PR.10 | The Product, when stained using the HER2 BMA Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 12 | PR.11 | Following a minimum of 14 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. The Product must also retains its ability to correctly detect procedural failures. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 13 | PR.12 | The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide after being stored at room temperature for at least six months. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 14 | PR.13 | Three lots of the Product, manufactured from separate multiblocks will be stained using a single lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by one trained pathologist, with the lot of the Product as the only variable. Staining results must align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 15 | PR.14 | Sections taken from the beginning, middle, and end of a single Cell Line Control multiblock will be stained using one lot of the BOND Ready-to-Use Primary Antibody HER2 BMA. A trained pathologist will evaluate the slides, with the section position within the multiblock as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 16 | PR.15 | Cell Line Control slides from the same lot will be stained on five non-consecutive days (spanning a minimum period of 20 days). A single trained pathologist will score the slides, with the staining day being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 17 | PR.16 | Cell Line Control slides from the same lot will be stained using three different BOND-III fully automated systems. A single trained pathologist will score the slides, with the BOND-III instrument being the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 18 | PR.17 | Cell Line Control slides from the same lot will be stained on different BOND Runs/Slide Staining Assemblies (SSAs) within the same BOND Run Batch. A single trained pathologist will score the slides, with the BOND Run/SSA as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 19 | PR.18 | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Polymer Refine Detection (DS9800) system. The slides will be scored by a single trained pathologist, with the lot of the DS9800 as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 20 | PR.19 | Cell Line Control slides from the same lot will be stained using three individual lots of the BOND Ready-to-Use Primary Antibody HER2 BMA. The slides will be scored by a single trained pathologist, with the lot of the antibody as the only variable. Staining results must meet the scoring criteria outlined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 21 | PR.20 | Cell Line Control slides from BOND runs performed at environmental temperatures between 18-26°C must stain as expected - Core A (BT474-M3) must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 22 | PR.21 | Cell Line Control slides from BOND runs performed at environmental humidities between 30-80% Relative Humidity must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 23 | PR.22 | HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA Assay Interpretation Guide. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 24 | PR.23 | Users will review the Product slide mailer label, hazard notification slip and IFU and complete a checklist to confirm that the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents) is present. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 25 | PR.24 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 26 | PR.25 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 27 | PR.26 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 28 | PR.27 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 29 | PR.28 | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 30 | PR.29 | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remain fully adhered and undamaged. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 31 | PR.30 | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 32 | PR.31 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 33 | PR.32 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 34 | PR.33 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 35 | PR.34 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 36 | PR.35 | The final labelling for the product (slide mailer and shipping labels, hazard notification slip and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 37 | PR.36 | Cell Line Control slides from BOND runs performed at altitudes up to and including 1600 meters above sea level must produce staining results that align with the scoring criteria defined in the HER2 BMA Interpretation Guide. Sample Size: TBC | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 38 | PR.37 | Users will review the Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 39 | PR.38 | All Cell Line Controls must maintain acceptable staining and cell morphology when slides are dehydrated using xylene and alcohol following completion of staining on the BOND-III using the HER2 BMA Assay System. All Cell Line Controls must maintain acceptable staining and cell morphology when slides are coverslipped using either xylene or toluene based mounting media following completion of staining on the BOND-III using the HER2 BMA Assay System. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 40 | PR.39 | The HER2 BMA Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA staining run. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP61-VEP-001 | 5 | 41 | PR.40 | The SDS must fully comply with all applicable regulatory standards, including: OSHA Hazard Communication Standard (29 CFR 1910.1200), REACH Regulation (EC No 1907/2006), CLP Regulation (EC No 1272/2008), GHS formatting and classification requirements. The SDS must include all 16 sections as mandated by GHS: Identification, Hazard(s) Identification, Composition/Information on Ingredients, First-Aid Measures, Fire-Fighting Measures, Accidental Release Measures, Handling and Storage, Exposure Controls/Personal Protection, Physical and Chemical Properties, Stability and Reactivity, Toxicological Information, Ecological Information, Disposal Considerations, Transport Information, Regulatory Information, Other Information. The final SDS must be reviewed and approved by an associate trained in SDS generation. The reviewer must confirm that the SDS contains accurate, complete, and up-to-date safety information. Sample Size: TBD | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP52.DP61-VAP-002 | 5 | 2 | N/A – Marketing Requirement | Performance will be deemed acceptable if the agreement rate between HER2 FISH status and HER2 BMA CDx IHC status on discordant cases (HER2 BMA CDx Assay vs. Comparator) is ≥90% for PPA and NPA. Respective lower bounds for each will be reported. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52.DP61-VAP-002 | 6 | 3 |  | BOND Ready-to-Use Primary Antibody HER2 BMA CDx will be used on intended use tissue with no failures attributed to tissue type. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52.DP61-VAP-002 | 6 | 4 |  | No safety issues or adverse events raised by site personnel during external performance studies in relation to The Product. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52.DP61-VAP-002 | 7 | 3 |  | 100% of evaluated HER2 BMA Cell Line Control slides must contain three cores with expected HER2 expression levels. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52.DP61-VAP-002 | 7 | 4 |  | The HER2 BMA Cell Line Control shall demonstrate distinguishable null, low, and high HER2 staining characteristics following processing with the HER2 BMA CDx Assay on the BOND-III instrument and provide interpretable semiquantitative evidence of staining performance in 100% of evaluated slides on runs deemed valid. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52.DP61-VAP-002 | 7 | 5 |  | The HER2 BMA Cell Line Control when stained with the HER2 BMA antibody demonstrates distinguishable null, low, and high HER2 staining characteristics that can be read by a qualified pathologist to demonstrate run validity in 100% of evaluated slides as per the HER2 BMA CDx Assay interpretation guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52-VEP-003 | 10 | 2 | 1 | All checklist items pass | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 10 | 3 | 2 | All checklist items pass | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 10 | 4 | 3 | All checklist items pass | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 10 | 5 | 4 | All checklist items pass (Where applicable) | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 10 | 6 | 5 | All checklist items pass | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 3 | dRMW-016 Risk Lines 001, 007, 010 and 016 LBN-euRMW-020 Risk Lines 001, 003 009, 011 and 013 | Instructions For Use (IFU) document specifies use of The Product with formalin fixed, paraffin-embedded (FFPE) breast carcinoma tissue prepared using the recommended pre-analytic conditions (fixative type, fixation time, processing parameters and sectioning thickness). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 4 | LBN-dRMW-016 Risk Line 001, 007, 010 and 016 LBN euRMW-020 Risk Line 001 | Product Instructions For Use (IFU) document contains details of any limitations related to tissue fixation (e.g. fixative type, fixation duration, over or under fixation effects). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 5 | LBN-euRMW-020 Risk Lines 055 and 093 | Product Instructions For Use (IFU) document specifies the recommended storage conditions for cut tissue sections during the stability period (temperature, humidity, protection from light). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 6 | LBN-dRMW-016 Risk Line 034 LBN-euRMW-020 Risk Lines 078, 089 and 097 | Product Instructions For Use (IFU) document contains sequential list of steps of the recommended procedures, including an immunohistochemistry (IHC) staining protocol, heat-induced epitope retrieval (HIER) protocol parameters, any pre-analytical recommendations, controls to be used, and scoring guidance. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 7 | LBN-dRMW-016 Risk Lines 086, 088 and 089 LBN euRMW-020 Risk Line 095 | Product Instructions For Use (IFU) states any limitations on product use related to altitude, temperature, and humidity. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 8 | LBN-dRMW-016 Risk Lines 071, 072, 073, 074 and 075 LBN-euRMW-020 Risk Line 044, 046 and 047 | Product reagent container label, hazard notification slip and IFU includes the required safety information (defined on the completed Form6030 - CLP Hazard Classification Of Chemical Reagents). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 9 | LBN-dRMW-016 Risk Lines 078 and 081 | All symbols and text included on the product labelling must be human legible in accordance with all applicable regulatory standards for USA, Canada, EU, UK. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 10 | LBN-dRMW-016 Risk Line 068 LBN-euRMW-020 Risk Lines 030, 033.5, 039, 044, 046,047, 091 and 095 | Laboratory users will confirm that the hazards detailed in the SDS can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP), as defined in the completed Form 6030 (CLP Hazard Classification of Chemical Reagents) and COSHH assessment. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 11 | N/A | The final labelling for the product (container and shipping labels, hazard notification slip, IFU and SDS) are compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 12 | LBN-dRMW-016 Risk Line 071 | The Product (container and shipping labels, hazard notification slip, IFU and SDS) must contain the relevant elements as per US regulation 21 CFR Part 801 and 809, "Labelling for in vitro diagnostic products". | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 13 | LBN-dRMW-016 Risk Line 071 | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 14 | LBN-dRMW-016 Risk Line 071 | Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 15 | LBN-dRMW-016 Risk Line 071 | The Product (container and shipping labels, hazard notification slip, IFU and SDS) contains the relevant elements as per UK general IVD labelling regulations MDR 2002. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 16 | LBN-dRMW-016 Risk Lines 010, 011, 016, 028, 039 and 040 LBN-euRMW-020 Risk Lines 019, 019.5, 020, 021, 022.1, 023 | HER2 BMA CDx Assay training materials and Proficiency Test must be available to the user. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 17 | LBN-dRMW-016 Risk Lines 001, 007, 010, 016, 019, 020.5, 023, 024, 025, 028, 029, 030.1, 034, 042, 043, 046, 050, 053, 058, 059, 060, 062, 066, 068, 069, 084, 086, 087, 088 and 089 LBN-euRMW-020 Risks Lines 001, 003, 005, 009, 011, 011.5, 013, 015, 017, 044, 046, 047, 048, 049, 050, 052, 053, 055, 058, 064, 066, 068, 070, 072, 075, 078, 080, 083, 086, 089, 091, 092, 093, 095, 096, 097 and 098 | The Product Instruction for Use (IFU) and HER2 BMA Assay Interpretation Guide include requirements listed in the Risk Control Measures section of the Project Swift Risk Management Worksheets. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 12 | 18 | LBN-dRMW-016 Risk Lines 068, 071, 081 and 084 LBN-euRMW-020 Risks Lines 044 | The SDS and Hazard notification must fully comply with all applicable regulatory standards for USA, Canada, EU, UK. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 55 | dRMW-016 | Background/non specific staining present and cell lines void the run. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 56 | dRMW-019 | Laboratory unable to achieve the desired number of patient samples to be stained, and the lab does not have additional stock of antibody units. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 57 | dRMW-020.5 | Insufficient volume dispensed on slide leading to poor staining quality. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 58 | dRMW-023 | The product is loaded onto a non BOND III instrument and accepted. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 59 | dRMW-024 | The product is loaded onto a BOND III instrument and is accepted and run on a non validated version of BOND software. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 60 | dRMW-025 | The product stains inappropriately using current BOND III bulk reagents and ancillaries | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 61 | dRMW-028 | The Interpretation Guide is inadequate to facilitate HER2 scoring | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 62 | dRMW-029 | Tissue sections are stained sub-optimally when different lots of antibody are used | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 63 | dRMW-030.1 | Tissue sections are stained sub-optimally when different lots of detection kit are used | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 64 | dRMW-034 | The product stains sub-optimally using the recommended staining protocol | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 65 | dRMW-039 | Staining of the tissue is interpretated differently by different pathologists | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 66 | dRMW-040 | Staining of the tissue is interpretated differently by the same pathologist on a different day leading to inconsistent diagnosis | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 67 | dRMW-042 | The product results in sub-optimal staining when used on different BOND III instruments | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 68 | dRMW-043 | The product is loaded onto the BOND and results in sub-optimal staining between different SSAs within the same BOND-III instrument | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 69 | dRMW-046 | The product is loaded onto the BOND-III and results in sub-optimal staining when used on a different day | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 70 | dRMW-050 | Staining becomes sub-optimal over the life of the reagent due to dilution of reagents | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 71 | dRMW-053 | Staining becomes sub-optimal over the life of the reagent due to evaporation | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 72 | dRMW-058 | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 73 | dRMW-059 | The product does not have the minimum stability | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 74 | dRMW-060 | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 75 | dRMW-062 | The product is loaded onto the BOND-III instrument and tissue sections are stained sub-optimally | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 76 | dRMW-066 | The product requires mixing or shaking to produce optimal staining | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 77 | dRMW-068 | The product contains hazardous material that cannot be adequately controlled by good laboratory practices | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 78 | dRMW-069 | The product is loaded onto the BOND III instrument for 16 & 23 runs and is not stable for all cycles of use | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 79 | dRMW-070.5 | The product cannot be registered into the BOND inventory by scanning the side barcode | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 80 | dRMW-071 | The user is unable to determine the product and/or hazards as per their territory | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 81 | dRMW-072 | The product is opened and the user cannot read the labelling font or the symbols on the container labels so cannot identify the product | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 82 | dRMW-073 | The product is opened and the user cannot read the labelling font, symbols or barcode on the container labels so cannot identify the product | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 83 | dRMW-074 | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 84 | dRMW-075 | The product is received and the user cannot read the labelling font or the symbols on any of the product labelling | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 85 | dRMW-078 | The user cannot read the labelling font or the symbols on the container labels after exposure to product reagents or BOND reagents, rendering the user unable to identify relevant safety labelling | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 86 | dRMW-081 | After normal storage and handling the labels are not legible by the user, rendering the user unable to identify the reagent/relevant safety labelling | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 87 | dRMW-084 | The product leaks after transport, exposing the user to the antibody | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 88 | dRMW-086 | The product does not provide acceptable staining across the range of altitudes defined in the BOND user manual due to atmospheric pressure affecting the HIER steps | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 89 | dRMW-087 | The product is stained on BOND but is not compatible with laboratory clearing and mounting procedures which results in difficulty in interpretation | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 90 | dRMW-088 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of temperatures defined in the BOND-III user manual | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 91 | dRMW-089 | The HER2 BMA CDx Assay does not provide acceptable staining across the range of operating humidities defined in the BOND user manual | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 92 | dRMW-090 | Hazard notification slip may fall out and user won't be aware of hazards and know how to find IFU | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 93 | dRMW-091 | Hazard notification slip and products may be exposed to environmental condition affecting legibility of text and labels | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-003 | 13 | 94 | dRMW-093 | User selects an alternative non-recommended HIER protocol leading to sub-optimal staining. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-005 | 9 | 2 | Prevalence | Results obtained from staining with the HER2 BMA Antibody CDx must reflect real-world prevalence of each HER2 IHC score category. This links to the PR detailed in Table 7. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-005 PREVALENCE Core Team Review.docx
DP52-VEP-005 | 10 | 2 | dRMW-016 Risk Lines 001, 002, 003, 007, 010, 011 and 016 LBN-euRMW-020 Risk Lines 001, 003, 005, 009, 011 and 013TBC | Tissue sections stained with the Product must produce staining that covers all HER2 expression categories (IHC 3+, IHC 2+, IHC 1+, IHC 0 with or absent membrane staining) across the range seen in the general population. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-005 PREVALENCE Core Team Review.docx
DP52-VEP-007 | 16 | 2 | dRMW-risk-029 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance when different lots of BMA antibody are used | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%.Point estimate of ≥90% when the lower bound of the 95% Confident Interval is ≥85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52-VEP-007 | 16 | 3 | dRMW-risk-030.1 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance when different lots of detection kits are used | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52-VEP-007 | 16 | 4 | dRMW-risk-042 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance across different BOND III instruments (inter instrument) | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52-VEP-007 | 16 | 5 | dRMW-risk-043 (LBN-dRMW-016)2: The product does not maintain comparable and acceptable staining performance both inter staining runs on BOND-III (Inter Run) | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52-VEP-007 | 16 | 6 | dRMW-risk-043 (LBN-dRMW-016): The product does not maintain comparable and acceptable staining performance within BOND runs/SSAs (Intra run) | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52-VEP-007 | 16 | 7 | dRMW-risk-046 (LBN-dRMW-016): The product does not maintain comparable and acceptable performance when used across different days (inter day) | Positive Percent Agreement (PPA), Negative Percent Agreement (NPA), and Overall Percent Agreement (OPA) point estimates ≥90%, with the lower bound of the two‑sided 95% CI ≥85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52-VEP-008 | 10 | 2 | Tissue Thickness | The agreement rate for the staining intensity score compared to the control (4µm) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 10 | 3 | Slide Baking | The agreement rate for the staining intensity score compared to the control (standard baking condition) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 10 | 4 | Mounting Media | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 10 | 5 | Delayed Start (7-Days) | 12/12 agreement in staining intensity score compared to the control (No Delay). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 10 | 6 | Hydration Tolerance (24-hour) | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 10 | 7 | HIER Time, Temp, and ER | TBC – Align with Craig | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 12 | 2 | dRMW-016 Risk Lines 003, 020.5, 025, 029, 030.1, 042, 043, 046, 050, 053, 058, 059, 060, 062, 066 and 069 LBN-euRMW-020 Risk Line 011.5, 024 and 072 | 100% of valid runs using the HER2 BMA Cell Line Control (CLC) slides when stained with The Product on the BOND-III instrument, must score as expected (Core A (BT474-M3) must be scored as HER2 "High Expression", Core B (RT4) must be scored as HER2 "Low Expression" and Core C (PANC-1) must be scored as HER2 "No Expression") as defined in the HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 12 | 3 | euRMW-020 Risk Lines 093 | The HER2 status of tissue samples stained with the Product following a delay of seven days between placement on the BOND-III instrument and the start of the staining run, must be >90% concordant with the results obtained from staining performed immediately after placement. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 12 | 4 | euRMW-020 Risk Line 096 | The HER2 status of tissue sections stained with the Product and maintained in a hydrated state on the BOND-III instrument for up to 24 hours must be >90% concordant with the HER2 status of matched sections removed immediately after run completion. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 12 | 5 | LBN-dRMW-016 Risk Line 087 LBN-euRMW-020 Risks Line 053 | The agreement rate for the staining intensity score compared to the control (DPX Mounting Media) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-008 | 13 | 2 | euRMW-risk-011.1 | The agreement rate for the staining intensity score compared to the control (standard baking condition) shall be at least 85% as demonstrated by the lower bound of the two-sided 95% confidence interval. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-010 | 11 | 6 | Fixation testing (cold ishemiaischemia time) | The sample size is based on complete coverage of the predefined cold ischemia time conditions (0, 0.5, 1, 3, 6, 12 and 24 hours). Comparisons will be performed relative to the reference condition (0-hour delay to fixation), resulting in six pairwise comparisons. For each comparison, agreement with the reference HER2 classification will be recorded and summarised as the proportion of comparisons matching the reference classification. The study is descriptive in nature and is intended to characterize the impact of delayed fixation on HER2 staining performance rather than demonstrate assay performance through statistical hypothesis testing. The selected timepoints represent a range of clinically relevant pre-analytical conditions and enable assessment of the impact of increasing delay to fixation on staining performance.Six pairwise comparisons will be performed, with each time point (0.5h, 2h, 3h, 6h, 12h, 24h) compared to the 0h duration. For each comparison, agreement in HER2 classification (positive/negative) will be recorded. Agreement rates will be summarised, as the proportion of comparisons matching the reference classification. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-010 Sample Characterisation Ver Protocol.docx
DP52-VEP-010 | 12 | 2 | dRMW-016 Risk Lines 001, 002, 003, 007, 010, 011 and 016 LBN-euRMW-020 Risk Lines 001, 003, 005, 009, 011 and 013 | The Product must stain ≥90% of test parameter FFPE breast carcinoma tissue slides with the same HER2 score as the control parameter to acquire the associated claim (fixation type, fixation duration and cold ischemia testing). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-010 Sample Characterisation Ver Protocol.docx
DP52-VEP-011 | 11 | 4 | Tour of Tumour | Staining should be Cconsistent with the expected tissue specificity of HER2, with non-target tumour tissues demonstrating no staining or only staining patterns consistent with those defined in in DP52-SAC-001. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-011 | 11 | 5 | Western Blot | With one unique cell lysate, suitability will be determined by specific BMA binding to HER2, indicated by a band at the expected molecular weight. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-011 | 12 | 2 | dRMW-risk-001: Antibody is not specific/sensitive and does not detect HER2 in intended tissue dRMW-risk-002: Antibody is not specific and does not detect HER2 in intended tissue dRMW-risk-003: Antibody is not sensitive and does not detect HER2 in intended tissue dRMW-risk-007: Antibody is not specific/sensitive and does not detect HER2 in intended tissue dRMW-risk-010: Antibody is not specific and wrongly detects HER2 in intended tissue (background/non-specific staining present) dRMW-risk-011: Antibody is not specific and wrongly detects HER2 in intended tissue dRMW-risk-016: Antibody is not specific and wrongly detects HER2 in intended tissue. | WESTERN BLOT Band must align with the expected ladder band corresponding to HER2 (~185 kDa). PEPTIDE INHIBITION Tissue cases stained with the Primary Antibody product pre-incubated with specific blocking peptide shall have decreased or loss of specific staining at a range of dilutions when compared to same samples stained with the Primary Antibody only and the product incubated with the non-specific peptide. TOB/TOT ≥95% in Normal Tissue must meet the staining described in DP52-SAC-001. ≥95% in Abnormal Tissue must meet the staining described in DP52-SAC-001. TMAs (or whole tissue cases) scored as per PCS Work Instruction. All staining will be reported and should be consistent with that described in the literature and Scoring and Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-011 | 13 | 4 | Training | BOND Operators will be trained in NCL-WI-000185 (R&D), NCL-WI-000232 and NCL-WI-000213 which will be documented in IQVIA. Microtomists will be trained in NCL-MFG-WI-000162, which will be documented in IQVIA. Scorers who score the peptide inhibition slides will be trained in accordance to DP52/DP61-SIG-001, DP52/DP61-QRG-001 and DP52/DP61-PTD-001. Scorers who have passed the proficiency test will have an approved Training Record Form (NCL-FRM-000319) generated. Scorers who score the TOB/TOT will be a PCS SME. Documentation of all training and proficiency assessments will be maintained within the DDHF and e-training files on the T:;Drive. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-011 | 13 | 5 | Facilities | Sample preparation, testing and scoring will be conducted at LBS-NCL. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-011 | 13 | 6 | Equipment | One BOND-III instrument will be used to conduct this study. Software version 7.4 or above will be used to conduct this study BDD Version 1213 or above will be used to conduct this study. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-011 | 13 | 7 | Reagents | A single lot of PP8680 manufactured as per the MDF. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-012 | 7 | 2 | Real Time Stability | 100% of all HER2 test slides stained with three lots of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52-VEP-012 | 7 | 3 | Transport Stability | 100% of all HER2 test slides stained with one lot of The Product that has been subjected to routine transport conditions must have an equivalent HER2 score to their respective Day 0 baseline (pre-transport) slides to acquire the associated reagent stability transport claim | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52-VEP-012 | 7 | 4 | In Use Stability | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. The Product must obtain a minimum shelf-life of six months with a two-month safety margin. | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52-VEP-012 | 8 | 2 | UN.04 | 100% of all HER2 test slides stained with one lot of The Product (aged at 2-8°C) must have an equivalent HER2 status to their respective Day 0 baseline (pre-transport) slides following thermal cycling at each specified timepoint to acquire the associated reagent stability claim. | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52-VEP-012 | 9 | 2 | Reagent immediate labelling verification | *Product container labels shall remain adhered to their container throughout the claimed product life and shall exhibit no lifting, peeling, detachment, or loss of adhesion up to and including the most recently passed stability timepoint used to support shelf-life claims. | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52-VEP-012 | 40 | 2 | 4 | No change in HER2 status (HER2-Null, HER2-UL, HER2-1+, HER2-2+ or HER2-3+) compared to baseline, as determined by DP52.DP61-SIG-001membrane staining completeness, staining intensity, and the percentage of invasive tumour cells exhibiting that pattern. . N.B: Changes obviously attributable to tissue heterogeneity or clear changes in morphology (e.g. cutting through) will not result in failure. These slides will be tested with bracketing. | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52-VEP-013 | 10 | 2 | Assay System Compatibility PR.02 | The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 10 | 3 | DMS setup test PR.12, PR.19, PR20 and PR.21 | See Table 8 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 10 | 4 | Ready to Use test PR.13 | Tissue cases stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, 100% of slides must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 10 | 5 | Labelling checks PR.45 | The Product when expired, will not allow a staining run to commence. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 10 | 6 | Reagent System Compatibility PR.39 | The Product must generate Acceptable Staining when used within the same BOND Run Batch as other breast BOND Ready-to-Use antibodies. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 10 | 7 | Reagent System Compatibility PR.40 | The Product can be run within the same BOND Run Batch as other BOND Detection systems and staining protocol combinations. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 10 | 8 | Control Tissue Compatibility PR.59 | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/Null) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 2 | LBN-dRMW-016 Risk Lines 023, 024, 025, 030.1 and 034 LBN-euRMW-020 Risk Lines 015, 017, 048, 049, 050, 052, 058, 064, 066 and 080dRMW-risk-025 | The Product must run correctly when loaded and run on the BOND-III instrument with all HER2 BMA Assay System components. 100% of BOND Run Events, if generated, must show no failure modes that are attributable to compatibility issues.The Product must be correctly loaded and run on the BOND-III instrument with all system components. BOND Run Events will be evaluated to ensure no failure modes are attributed to compatibility issues. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 3 | LBN-dRMW-016 Risk Line 023 and 024 LBN-euRMW-020 Risk Line 068 | The Product, when loaded onto the BOND-III instrument must be recognized by the BOND controller and must complete a staining run using the default protocols (IHC/HIER) assigned to the Product. The Product, when loaded onto the BOND RX, BOND RXm instruments, should not be recognized or registered and unable to initiate a staining run. When loaded onto a BOND-MAX or BOND-PRIME the User must be unable to initiate a staining run. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 4 | LBN-dRMW-016 Risk Lines 062 and 066 LBN-euRMW-020 Risk Line 092 | 100% of slides stained with The Product after removal from the recommended storage conditions (2–8 °C), and without any additional preparation, must demonstrate the expected staining characteristics as defined in the HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 5 | LBN-dRMW-016 Risk Lines 024, 070 and 070.5dRMW-risk-070 dRMW-risk-070.5 dRMW-risk-080 | When the barcode on the front of The Product reagent container label is scanned using the hand-held scanner connected to a BOND-III instrument, the system must recognize the reagent. The information displayed in the dialog box fields - "Name", "Catalogue No", "UPI", "Lot No", "Expiration Date", and "Initial vol. (mL)", must match the details printed on the reagent container label. When the barcode on the top of the product reagent container label is scanned by the on-board scanner of a BOND-III instrument, the abbreviated device name must be correctly displayed on the BOND interface. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 6 | LBN-dRMW-016 Risk Lines 034 LBN-euRMW-020 Risk Lines 066 and 078dRMW-risk-034 | When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated, and the instrument will not allow any changes to be made to the protocol settings.When The Product is scanned on to the BOND-III, the default IHC staining protocol will be pre-populated and the instrument will not allow any changes to be made to the protocol settings. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 7 | LBN-dRMW-016 Risk Line 086 LBN-euRMW-020 Risk Lines 001, 089 and 098dRMW-risk-086 dRMW-risk-034 dRMW-risk-088 dRMW-risk-089 dRMW-risk-093 | When The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so theWhen The Product is scanned on to the BOND-III, the default HIER protocol will be pre-populated but unlocked, so the instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. instrument will allow for a new HIER protocol to be selected, and EIER is locked to no protocol. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 8 | LBN-dRMW-016 Risk Lines 056 and 057dRMW-risk-056 dRMW-risk-057 | The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance.The Product IHC staining protocol must run to completion on the BOND-III instrument, whilst running alongside other breast BOND Ready-to-Use (RTU) antibody protocols, on separate SSAs within the same BOND Run Batch. Stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 9 | dRMW-risk-056 dRMW-risk-057 dRMW-risk-034N/A | The Product, when run with the BOND Polymer Refine Detection kit (DS9800) must run to completion on BOND-III instrument, whilst running alongside other BOND-III detections kits: BOND Polymer Refine Red Detection on (DS9390), ChromoPlex I Dual Detection (DS9665), and ChromoPlex II Dual Detection (DS9494), on separate SSAs within the same BOND Run Batch. 100% of stained tissue sections from all BOND runs must exhibit acceptable staining as defined in the individual device specific guidance. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 10 | LBN-dRMW-016 Risk Line 058 LBN-euRMW-020 Risk Line 042dRMW-risk-058 | Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument.Users must be unable to start a staining run when expired Product is loaded onto BOND-III instrument. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 11 | 11 | LBN-dRMW-016 Risk Lines 001, 007, 010, 016, 020.5, 025, 029, 030.1, 042, 043, 046, 050, 053, 058, 059, 060, 062, 066 and 069 LBN-euRMW-020 Risk Lines 001, 003, 011.5, 021, 022.1, 042, 075, 080, 083, 089, 091, 095, 097 and 098 | FFPE tissue suitable for use as positive/negative control tissue (breast carcinoma tissue displaying moderate tumor cell membrane staining - HER2 2+/HER2-Null respectively) must display expected staining patterns and intensity when stained with the HER2 BMA Assay. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP52-VEP-013 | 12 | 2 | EIER is locked to no protocol. | EIER is locked to no protocol to prevent use when scanning the BOND Ready-to-Use Primary Antibody HER2 BMA CDx. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP61-VEP-002 | 8 | 2 | Product Staining Interpretation | HER2 BMA CDx Assay Interpretation Guide includes representative images and clearly describes how staining of the Cell Line Control can be interpreted - checklist | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 3 | Cell Core Staining Interpretation | Confirm the HER2 BMA CDx Assay Interpretation Guide describes the expected staining patterns and characteristics of each cell line core - checklist | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 4 | Hazard and Safety Labeling | To confirm the Product has labelling which provides all required hazard and safety information and associated symbols | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 7 | Control of Hazards by PPE and GLP | To confirm the Product does not present any hazards to end users or the environment that cannot be adequately controlled by good laboratory practices (GLP) and the use of standard personal protective equipment (PPE) | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 9 | On-slide Label Legibility Following Transport Cycling, Storage and Use | To confirm the Product on-slide labels remain human legible after shipping and storage within the recommended temperature range and normal use. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 10 | Slide Mailer Label Attachment Following Transport Cycling, Storage and Use | To confirm the Immediate packaging (slide mailer) Label remains attached after shipping and storage within the recommended temperature range and normal use. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 11 | Slide Mailer Label Legibility Following Transport Cycling, Storage and Use | To confirm the symbols and text included on Product labelling are human legible following shipping and storage within the recommended temperature range and normal use. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 12 | Slide mailer serviceability following Transport Cycling, storage and use | Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 13 | UDI Specifications Compliance | A trained RA associate will confirm the final product labelling (slide mailer and IFU) is compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 15 | US regulation 21 CFR Part 801 and 809 Compliance | A trained RA associate will confirm the final product labelling (slide mailer and IFU) is compliant with US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 17 | IVDR EU/2017/746 Compliance | A trained RA associate will confirm the final product labelling (slide mailer, shipping labels and IFU) is compliant with Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 19 | Canadian IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23 Compliance | A trained RA associate will confirm the final product labelling (slide mailer, shipping labels and IFU) is compliant with Canadian IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 21 | UK general IVD labelling regulations MDR 2002 Compliance | A trained RA associate will confirm the final product labelling (slide mailer, shipping labels and IFU) is compliant with UK general IVD labelling regulations MDR 2002. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 8 | 23 | Project Swift Risk Management Worksheets Information | To confirm the appropriate content in the Recommended Risk Control Measures of the Project Swift Risk Management Worksheets is present in the Product Instruction For Use (IFU) and Interpretation Guide (IG) | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 2 | LBN-dRMW-017 Risk Line 013 LBN-euRMW-020 Risk Line 024 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 3 | LBN-dRMW-017 Risk Line 013 LBN-euRMW-020 Risk Line 024 | The HER2 BMA CDx Assay Interpretation Guide must be reviewed by scoring subject matter experts (SMEs) to ensure it includes: - Detailed, step-by-step instructions for interpreting the HER2 BMA Cell Line Control staining results. - Instructions clearly defining expected staining patterns for each cell line core. - High-resolution images that are representative of the staining patterns expected for each cell line core. Images must be annotated and accompanied by relevant interpretive notes. - Clear, objective pass and fail criteria for validating the HER2 BMA CDx staining run. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 4 | LBN-dRMW-017 Risk Line 028 | Users will review the Product slide mailer label, IFU and product specifications and complete a checklist to confirm that the required safety information is present. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 5 | LBN-dRMW-017 Risk Line 028 | Laboratory users will confirm that the hazards detailed in the IFU and product specifications can be adequately controlled under normal use conditions by employing standard personal protective equipment (PPE) including safety glasses, lab coat, and nitrile gloves and following Good Laboratory Practice (GLP). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 6 | LBN-dRMW-017 Risk Line 032, 032.5, 033 and 033.5 LBN-euRMW-020 Risk Lines 035 and 037 LBN-dRMW-017 Risk lines 032, 032.5, 033, 033.5 050, 051 LBN-euRMW-020 Risk Line 035 and 037 | The Product’s Immediate Packaging (slide mailer) must remain intact, undamaged, and safe to handle (no sharp edges) after exposure to conditions simulating routine shipment and storage. Following routine shipping, storage and use conditions, R&D lab users will inspect the slide mailer and complete a checklist to confirm that the slide mailer is still serviceable, and functions as intended. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 7 | LBN-dRMW-017 Risk Line 032, 032.5, 033 and 033.5 LBN-euRMW-020 Risk Lines and 035 | When packaged within the Outer Packaging (pouch), the Product must remain intact, undamaged, and safe to handle with no breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will inspect the Outer Packaging (pouch) for any damage (e.g., tears, punctures) and complete a checklist to confirm that the Product is intact and safe to handle. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 8 | LBN-euRMW-020 Risk Lines 027 and 035 | When packaged in the Shipping Packaging (cardboard box), the Product must remain intact, undamaged, and safe to handle without any breakages or sharp edges following exposure to conditions representative of routine shipment and storage. Users will complete a checklist to confirm that the Product is intact and safe to handle. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 9 | LBN-dRMW-017 Risk Lines 031 and 041 | Following routine shipping, storage and use conditions, users will inspect the Product on-slide label and complete a checklist to confirm that the text is readable with no smearing, fading, or distortion. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 10 | LBN-dRMW-017 Risk Lines 043 and 044 | Following routine shipping, storage and use conditions, users will inspect the Product immediate packaging (slide mailer) and complete a checklist to confirm that the label has remained fully adhered and undamaged. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 11 | LBN-dRMW-017 Risk Line 041 | Following routine shipping and storage conditions, users will inspect the Product slide mailer label and complete a checklist to confirm that the barcode, symbols, and text exhibit no smearing, fading, or distortion. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 12 | N/A | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to be compliant with UDI specifications as per EN ISO 18113-2 section 4.2, US FDA regulation 21 CFR Part 801 and IVDR Regulation (EU) 2017/746 Article 24. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 13 | LBN-dRMW-017 Risk Line 029 | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per US regulation 21 CFR Part 801 and 809, 'Labelling for in vitro diagnostic products'. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 14 | LBN-dRMW-017 Risk Line 029 | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Class C (companion diagnostic) requirements from IVDR EU/2017/746, ISO15223-1 and ISO18113. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 15 | LBN-dRMW-017 Risk Line 029 | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per Canadian Class III IVD labelling regulations SOR / 98-282, Sections 21, 22 and 23. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 16 | LBN-dRMW-017 Risk Line 029 | The final labelling for the product (slide mailer and IFU) must be checked by an RA representative and confirmed to contain the relevant elements as per UK general IVD labelling regulations MDR 2002. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-002 | 9 | 17 | LBN-dRMW-017 Risk Lines 006, 011, 013, 028, 033, 033.5 and 049 LBN-euRMW-020 Risk Line 072, 075, 091 and 095 | Users will review the Product Instruction for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide and complete a checklist to confirm that all content requirements listed in the "Recommended Risk Control Measures" section of the Project Swift Risk Management Worksheets are identified and documented. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-003 | 9 | 2 | Inter-lot repeatability | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 9 | 3 | Intra-lot repeatability | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 9 | 4 | Inter-day reproducibility | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 9 | 5 | Inter-instrument repeatability | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 9 | 6 | Inter-run repeatability | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 9 | 7 | Intra-run repeatability | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 9 | 8 | Inter-detection lot repeatability | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 9 | 9 | Inter-antibody lot repeatability | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 10 | 2 | LBN-dRMW-017 Risk Lines 021 and 039 | Stained CLCs must demonstrate Acceptable Staining and Consistent Staining across six different lots. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.For inter-lot testing, the lower bound of the two-sided 95% confidence interval for the overall percent agreement should be at least 85%. Testing will be performed on 6 lots. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 10 | 3 | LBN-dRMW-017 Risk Line 022 | Stained CLCs For intra-lot testing, taken from within a single lot must generate Acceptable Staining and Consistent Staining. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate the overall percent agreement should be at least 85%. Only sections within a single lot will be tested. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 10 | 4 | LBN-dRMW-017 Risk Line 022 | When stained over 5 non-consecutive days (spanning a minimum period of 20 days), CLCs must generate Acceptable Staining and Consistent Staining. For inter-day testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate for the overall percent agreement should should be at least 85%. Testing will occur over 5 non-consecutive days spanning a minimum period of 20 days. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 10 | 5 | LBN-dRMW-017 Risk Line 022 | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across three different BOND-IIIs. For inter-instrument testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed across three BOND-IIIs. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 10 | 6 | N/A | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested across different SSAs on a BOND-III Instrument run batch and within BOND Runs (SSAs). For inter-run testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Iner-run data will be collected across BOND run batches. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 10 | 7 | LBN-dRMW-017 Risk Line 021.2 | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three different lots of BOND Polymer Refine Detection (DS9800) on the BOND-III system For inter-detectiontested samples lot repeatability testing, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Three lots will be used during testing. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-003 | 10 | 8 | LBN-dRMW-017 Risk Line 021.1 | Stained CLCs must generate Acceptable Staining and Consistent Staining when tested using three lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. For inter-antibody lot repeatability testingtested samples, the lower bound of the two-sided 95% confidence interval for the overall pass ratefor the overall percent agreement should be at least 85%. Testing will be performed using 3 lots of BOND Ready-to-Use Primary Antibody HER2 BMA CDx. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-005 | 11 | 2 | Product performance post transport simulation | 100% of stained Product must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide after being subject to transport simulation. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The pass rate OPA shall be ≥875% at demonstrated by the lower bound of the with atwo-sided 95% CI (confidence interval). The overall expected pass rate is ≥98%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 11 | 3 | Product shelf life stability | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The pass rateOPA shall be ≥875% as demonstrated by at the lower bound with aof the two-sided 95% CI. The overall expected pass rate is ≥98%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 11 | 4 | Intermediate product stability | Following a minimum of six months storage at room temperature, 100% of The Intermediate Product (individual agarose cell line cores of BT474-M3, RT4 and PANC-1) continues to produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The OPAoverall pass rate shall be ≥875% as demonstrated by at the lower bound with aof the two-sided 95% CI. The overall expected pass rate is ≥98%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 12 | 2 | LBN-dRMW-017 Risk Line 018 | 100% of stained Product must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide after being subject to transport simulation.The Product, after being subject to transport simulation, must meet the scoring criteria outlined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 12 | 3 | LBN-dRMW-017 Risk Line 017 LBN-euRMW-020 Risk Line 043 | Following a minimum of eight months storage at recommended storage conditions (2-8°C), 100% of The Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.Following a minimum of 18 months storage at recommended storage conditions, the Product continues to produce positive staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. The Product must also retain its ability to correctly detect procedural failures. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 12 | 4 | LBN-dRMW-017 Risk Line 039 | Following a minimum of six months storage at room temperature, 100% of The Intermediate Product (individual agarose cell line cores of BT474-M3, RT4 and PANC-1) continues to produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide.The Intermediate Product, consisting of FFPE blocks of individual cell line cores - BT474-M3, RT4 and PANC-1 must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide after being stored at room temperature for at least six months. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 9 | Ambient Temperature | 15 - 25°C (as room temperature definition) | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 10 | Baseline Reference | Comparator slides stained using suitable reference reagent at the baseline test point, used to assess whether staining performance at subsequent stability timepoints remains acceptable. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 11 | BDD | BOND Definition Database | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 12 | BMA | Breast Marker Antigen HER2 mouse monoclonal antibody clone | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 13 | BOND | The BOND system provided by Leica Biosystems. The system includes a range of reagents used for ISH and IHC staining, a robotic instrument for automated staining, and software to control the process. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 14 | BOND Run | A minimum of 1 and a maximum of 10 slides that are subjected to the automated staining process on a BOND-III fully automated system whilst positioned within the same slide tray and across the same Slide Staining Assembly (SSA) | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 15 | BOND-III | Automated robotic staining instrument, third‑generation BOND platform | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 16 | BOND-III fully automated system | BOND-III linked to a BOND Controller Computer equipped with BOND Software Version 7.4 or above, appropriate BDD Version (TBD123), hand-held barcode scanner and label printer | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 17 | BRTU | BOND Ready-to-Use Primary Antibody | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 18 | BSI | British Standards Institution | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 19 | CDx | Companion Diagnostic | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 20 | CI | Confidence Interval | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 21 | CLC | Cell Line Control | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 22 | CLSI | Clinical & Laboratory Standards Institute | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 23 | Day 0 | The day on which test units are placed into their assigned storage condition and aging begins, the day from which the shelf life length is based on. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 24 | DDP | Design and Development Plan | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 25 | DPX | Distyrene, Plasticizer and Xylene | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 26 | dRMW | Design Risk Management Worksheet | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 27 | eDHFeDDF | Electronic Design and Development FileDesign History File | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 28 | euRMW | End User Risk Management Worksheet | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 29 | FDA | U.S. Food and Drug Administration | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 30 | FFPE | Formalin Fixed Paraffin Embedded | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 31 | Freeze–Thaw Testing | Testing designed to evaluate the impact on reagent performance after one or more cycles of freezing and thawing, intended to simulate potential accidental exposure to sub‑zero temperatures during transport or handling. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 32 | FRM | Form | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 33 | HER2 | Human Epidermal Growth Factor Receptor 2 | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 34 | HER2 BMA CDx Assay | Consists of the BOND Ready-to-Use Primary Antibody HER2 BMA CDx and HER2 BMA Cell Line Control | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 35 | HER2 BMA CDx Assay System. | The HER2 BMA CDx Assay, in conjunction with the BOND-III fully automated system, BOND Polymer Refine Detection (DS9800), BOND Ready-to-Use Negative Control Negative (Mouse) (PA0996/PA0996-U), BOND ancillaries (BOND Dewax Solution (AR9222), BOND Wash Solution 10X Concentrate (AR9590), BOND Epitope Retrieval Solution 1 (AR9961), consumables BOND Universal Covertiles (S21.4611), BOND Plus Slides (S21.2113.A) or Apex BOND Slides (3800040), Instructions For Use (IFU) and HER2 BMA CDx Assay Interpretation Guide (SIG) are collectively referred to as the HER2 BMA CDx Assay System. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 36 | HIER | Heat Induced Epitope Retrieval | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 37 | IHC | Immunohistochemistry | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 38 | Immediate packaging | Slide mailer | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 39 | IDA | Industrial Denatured Alcohol | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 40 | LBS-NCL | Leica Biosystems Newcastle | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 41 | Lot | Defined amount of material, which is uniform in its properties and has been produced in one process or a series of processes. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 42 | MS | Manufacturing Science | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 43 | Multiblock | A paraffin wax block that contains three individual FFPE cell line cores and an agarose orientation core. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 44 | NPD | New Product Development | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 45 | NPI | New Product Introduction | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 46 | PPE | Personal Protective Equipment | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 47 | PTD | Pathologists Training Deck | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 48 | Positive/Negative Tissue Control | The control slide contains both HER2‑negative tissue (HER2‑Null) and HER2‑positive tissue (primarily HER2‑2+). The negative region confirms assay specificity, while the positive region verifies assay sensitivity and demonstrates weak to moderate complete tumor cell membrane staining.the expected staining intensity for a HER2 2+ control. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 49 | QA | Quality Assurance | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 50 | QC | Quality Control | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 51 | QRG | Quick Reference Guide | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 52 | R&D | Research & Development | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 53 | Real-Time Stability Testing | Stability study testing designed to examine the long-term robustness and stability of a product when exposed to conditions that are anticipated during transport, storage and use. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 54 | Real-Time Shelf Life Test | Testing designed to examine the long-term robustness and stability of a product when exposed to recommended storage conditions. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 55 | Room Temperature | 15 - 25°C Room temperature as it relates to testing has been defined as 27°C (+/- 1°C) to simulate the upper limit of assay operability temperature conditions for a laboratory environment containing a BOND-III Automated IHC and ISH Stainer (as per BOND User Manual version 78). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 56 | Shelf Life | Period of time until the expiry date, during which a product, in its original packaging, maintains its stability under the storage conditions specified by the manufacturer. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 57 | SIG | Scoring and Interpretation Guide | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 58 | SME | Subject Matter Expert | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 59 | SOP | Standard Operating Procedure | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 60 | SSA | An individual staining position within the BOND-III fully automated system, referred to as a Slide Staining Assembly (SSA). Each SSA can accommodate up to 10 slides during a staining run.Slide Staining Assembly | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 61 | SSR | Stability Summary Report | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 62 | Stability | The ability of a product, when kept under specified conditions, to retain its properties and/or performance throughout the specified shelf life. Stability assessment shall be based on data derived from real-time stability testing. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 63 | TO | Technical Operations | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 64 | Transport Simulation | Testing, after thermal cycling designed to simulate conditions that product might conceivably experience during transportation at ambient temperature. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-005 | 33 | 65 | WI | Work Instruction | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-006 | 9 | 2 | HER2 BMA CDx Assay System Compatibility | When stained using all components of the HER2 BMA CDx Assay system, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 9 | 3 | Ready to Use CompatibilityBOND-III Fully Automated System Compatibility | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 9 | 4 | Recommended HIER and IHC Staining Protocol CompatibilityReady to Use Compatibility | When stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 9 | 5 | Final Product Core Orientation ConsistencyRecommended HIER and IHC Staining Protocol Compatibility | A HER2 scoring SME/qualified pathologist will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product IFU and HER2 BMA CDx Assay Interpretation Guide. 100% of assessed CLC slides must demonstrate correct core orientation. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 9 | 6 | Final Product Core Orientation Consistency | The overall pass rate shall be ≥85% as demonstrated by the lower bound of the two-sided 95% confidence interval (CI). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 10 | 2 | LBN-dRMW-017 Risk Line 009 | When stained using all components of the HER2 BMA CDx Assay system, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using all components of the HER2 BMA CDx Assay system, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 10 | 3 | LBN-dRMW-017 Risk Line 009 | The Product, when loaded onto the BOND-III instrument, must not trigger any warning or error messages upon initiation of a staining run. The Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 10 | 4 | LBN-dRMW-017 Risk Lines 001 and , 004, 019 and 026 | The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). When stained within an hour of removal from 2–8°C storage, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product must be usable directly from 2–8°C storage without requiring any preparation (e.g., warming, baking, or incubation). The Product, when stained within an hour of removal from 2–8°C storage, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 10 | 5 | LBN-dRMW-017 Risk Line 026N/A | When stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, the Product must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.The Product, when stained using the HER2 BMA CDx Assay's recommended Heat Induced Epitope Retrieval (HIER) and IHC staining protocols, must produce staining results that align with the scoring criteria defined in the HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 10 | 6 | LBN-dRMW-017 Risk Line 006 | A HER2 scoring SME/qualified pathologist will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product IFU and HER2 BMA CDx Assay Interpretation Guide. 100% of assessed CLC slides must demonstrate correct core orientation. For tested samples, the lower bound of the two-sided 95% confidence interval for the overall pass rate should be at least 85%.HER2 scoring SMEs/qualified pathologists will review the slides to ensure the cores are consistently orientated relative to each other as described in the Product Instructions for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 12 | 6 | 4 | HER2 BMA Cell Line Control slides stained during Precision testing will utilize all components of the HER2 BMA CDx Assay system and will therefore be suitable to test against PR.04. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 12 | 7 | 5 | Precision testing of HER2 BMA Cell Line Control slides will only be performed on the BOND-III instrument and therefore data generated will be usable to verify PR.05. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 12 | 8 | 56 | SAll slides tested during the Precision study will be used directly from 2–8°C storage without any further preparation and within an hour of removal will be used to verify PR.09.. The movement of verification test materials into and out of cold storage shall be recorded in NCL-FRM-001244 to provide traceability and ensure the This will allow Precision testing data to be used to verify PR.09.can be used. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 12 | 9 | 67 | The recommended HER2 BMA CDx Assay HIER and IHC staining protocols will be used for all HER2 BMA Cell Line Control Precision testing allowing its use to support verification of PR.10. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 12 | 10 | 78 | All HER2 BMA Cell Line Control slides stained during Precision testing will be scored by HER 2 scoring SMEs/qualified pathologists and any deviation from the expected core orientation , as described in the Product Instructions for Use (IFU) and HER2 BMA CDx Assay Interpretation Guide, will be noted and used to verify PR.22. | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-006 | 12 | 11 | 89 | Results will be further detailed in the System Compatibility Verification Report for the HER2 BMA Cell Line Control (DP61‑VER-006). | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx

## Source Index
DocumentID | Revision | DocumentKind | AuthorityDecision | SourceFile
LBN-dRMW-016 | 7 | Risk workbook | DP52-specific design risk source | C:\Users\emma.goodrick\Downloads\OneDrive_1_22-09-2026\dRMW-016 Rev07 HER2 BMA CDx Primary Antibody - DRAFT.xlsx
LBN-dRMW-017 | 6 | Risk workbook | DP61-specific design risk source | C:\Users\emma.goodrick\Downloads\OneDrive_1_22-09-2026\dRMW-017 Rev06 HER2 BMA Cell Line Control - DRAFT.xlsx
LBN-euRMW-020 | 7 | Risk workbook | Shared DP52/DP61 end-user risk source | C:\Users\emma.goodrick\Downloads\OneDrive_1_22-09-2026\euRMW-020 Rev07 HER2 BMA CDx Assay - DRAFT.xlsx
DP52-PR-001 | 6 | Product Requirements | Signed source | C:\Users\emma.goodrick\Downloads\DP52-PR-001_Rev06_-_Product_Requirements_Signed 10-Aug-2026.pdf
DP61-PR-001 | 5 | Product Requirements | Signed source | C:\Users\emma.goodrick\Downloads\DP61-PR-001_Rev05_-_Product_Requirements_Signed 10-Aug-2026.pdf
DP52-VEP-001 | 1 | Overarching Plan | Signed source | C:\Users\emma.goodrick\Downloads\DP52-VEP-001_Overarching_Verification_Plan_for_BOND_Ready-to-Use_Primary_Antibody Signed 03-Aug-2026.pdf
DP52/DP61-VEP-001 (Environmental) | 1 | Protocol | Signed joint DP52/DP61 environmental protocol | C:\Users\emma.goodrick\Downloads\DP5261-VEP-001_-_Environmental_Verification_Protocol_for_the_HER2_BMA_CDx_Assay Signed 28-July-2026.pdf
RP101-IUVP-001 | 1 | Overarching Plan | Applied to DP52 PR.14 and PR.15 by signed DP52-VEP-001 | T:\R&D\PROJECTS\NPD Overarching Documents\In-Use Verification\RP101-IUVP-001. NPD Overarching In-Use Verification plan.pdf
RP101-IUVR-001 | 1 | Overarching Report | Completed evidence applied to DP52 PR.14 and PR.15 by signed DP52-VEP-001 | T:\R&D\PROJECTS\NPD Overarching Documents\In-Use Testing\RP101-IUVR-001_NPD_Overarching_In-Use_Verification_Report.pdf
DP61-VEP-001 | [NEEDS INPUT] | Overarching Plan | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-001 HER2 BMA Cell Line Control Overarching Verification Plan.docx
DP52.DP61-VAP-002 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52.DP61-VAP-002 Clinical Data Assessment NCL-FRM-002064 V&V Protocol.docx
DP52-VEP-003 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52‑VEP-003 Labelling Verification Protocol for the BOND Ready-to-Use Primary Antibody HER2 BMA CDx R&D Review Draft.docx
DP52-VEP-005 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-005 PREVALENCE Core Team Review.docx
DP52-VEP-007 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-007 Precision Core Team.docx
DP52-VEP-008 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-008 Robustness CORE TEAM REVIEW.docx
DP52-VEP-010 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-010 Sample Characterisation Ver Protocol.docx
DP52-VEP-011 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-011 Analytical Specificity and Sensitivity Draft1.docx
DP52-VEP-012 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\DP52-VEP-012 Reagent Stability Protocol Draft Core Team Review.docx
DP52-VEP-013 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP52-VEP-013 System Compatibility - Core Team Review.docx
DP61-VEP-002 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-002 Labeling Verification Protocol for the HER2 BMA Cell Line Control DRAFT SRH edit.docx
DP61-VEP-003 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-003 Precision Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx
DP61-VEP-005 | Draft supplied | Protocol | Authoritative UPDATE copy selected by user | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-005 Stability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team UPDATE.docx
DP61-VEP-006 | Draft supplied | Protocol | Approved source list | C:\Users\emma.goodrick\Downloads\OneDrive_2026-09-22\Individual Protocols\DP61-VEP-006 System Compatability Verification Protocol for the HER2 BMA Cell Line Control DRAFT Core Team.docx